## Frontpage
Author: | Stef Mestrom
R&D: Documentation Officer
Department: | Research and Development
Approval
Role
GEN: Project Manager
Q&R: Dev QA Officer
R&D: Department Manager IFU |  | (approval via electronic signature in document management tool)

## Revision History
Revision History
Revision Date | Author | Source | Changes/Comments
2022 Jan 25 | Frits Jacobs | see Note | Initial version
2022 Jul 12 | Frits Jacobs | see Note | Updated after review, components added
2022 Aug 23 | Frits Jacobs | see Note | Updated: • The value has been updated for L-arc-N 15 Azurion R3.0 systems (CVPRJ00419435) in: - Geometry - Frontal C-arm - Detector Rotation - Geometry - Frontal C-arm - Z-rotation - Geometry - Lateral L-arc - Longitudinal movement - Geometry - Lateral L-arc - Longitudinal movement speed - Geometry - Lateral L-arc - Detector movement - Geometry - Lareral L-arc - Detector movement speed • 3 new lines added for Assembly of X-ray tube and collimator (X-ray source assembly)
2022 Nov 18 | Stef Mestrom | see Note | Updated: Based on review comments it's f:ound out the minus sign when using RAO or LAO should be removed. Based on check we removed minus sign from - Geometry - Lateral L-arc - L-arc rotation Update Focal spot size and loadability to make in clearer 3.2 X-ray tube - Focal spot size and loadability
2023 jan 02 | Stef Mestrom | See Note | Updated: MRC 200 Tube information added MRC200+ Tube information updated TPO updated for Tubes Updated Video Output connection - Supported resolutions
2023 March 17 | Stef Mestrom | see Note | Updated: Wireless footswitch information Removed Examination Light information as it's captured in separate IFU of Examination light Removed need to include monitor data in IFU
2023 Dec 05 | Stef Mestrom | see Note | Updated: Included ref Air Kerma Tabs based on CVPRJ00437520 Azu.R3.x: Published Referenced AirKermaRate for Cardio Pediatrics Clarity > 40 kg not correct Field size tab Dose ratio FD20
Note: see the main IFU Specification DHF373049
Open Issues
None.

## Technical Data
Item | 3F12 | 3F15 | 5C12 | 5C20 | 5F20 | 7C12 | 7F12 | 7C20 | 7F20 | B12-12 | B20-12 | B20-15 | IFU | subitem | parameter | value | Ref Azurion 3.0 | Tag | TPO
x |  | x |  |  | x |  |  |  | x |  | x | x
Geometry | - | x | - | - | x | - | - | - | x | - | x | x | 1 | Frontal C-arm | Floor-Mounted Stand with 15-Inch or 20-Inch Detector | Tag: US.TechData.FrontalC-Arm.Clea_Floor | - | - | -
Geometry | - | x | - | - | x | - | - | - | x | - | x | x | 1 | Frontal C-arm | Iso-center to floor | 114.5 cm (45.1 inch) | DHF357242 | US.Pos.Clea.MainDimensionsClea - Figure 4 | Maikel van Eekelen
Geometry | - | x | - | - | x | - | - | - | x | - | x | x | 1 | Frontal C-arm | C-arm rotation | In head-end position: 120° LAO, 185° RAO, in side position: 90° LAO, 90° RAO | DHF357242 | US.Pos.Clea.MovementRotationAngulation - Stroke | Maikel van Eekelen
Geometry | - | x | - | - | x | - | - | - | x | - | x | x | 1 | Frontal C-arm | C-arm rotation speed | 25°/sec. | DHF357242 | US.Pos.Clea.MovementRotationAngulation - Motorized speed | Maikel van Eekelen
Geometry | - | x | - | - | x | - | - | - | x | - | x | x | 1 | Frontal C-arm | C-arm angulation | In head-end position: 90° cranial, 90° caudal In side position: 185° cranial, 120° caudal | DHF357242 | US.Pos.Clea.MovementRotationAngulation - Stroke | Maikel van Eekelen
Geometry | - | x | - | - | x | - | - | - | x | - | x | x | 1 | Frontal C-arm | C-arm angulation speed | 25°/sec. | DHF357242 | US.Pos.Clea.MovementRotationAngulation - Motorized speed | Maikel van Eekelen
Geometry | - | x | - | - | x | - | - | - | x | - | x | x | 1 | Frontal C-arm | Focal spot to iso-center | 81 cm (31.9 inch) | DHF357242 | US.Pos.Clea.MainDimensionsClea - Figure 4 | Maikel van Eekelen
Geometry | - | x | - | - | x | - | - | - | x | - | x | x | 1 | Frontal C-arm | Source Image Distance | 89.5-119.5 cm (35.2-47.0 inch) | DHF357242 | US.Pos.Clea.MainDimensionsClea - Figure 4 | Maikel van Eekelen
Geometry | - | x | - | - | x | - | - | - | x | - | x | x | 1 | Frontal C-arm | C-arm depth | 90 cm (35.4 inch) | DHF357242 | US.Pos.Clea.MainDimensionsClea - Figure 4 | Maikel van Eekelen
Geometry | - | x | - | - | x | - | - | - | x | - | x | x | 1 | Frontal C-arm | Detector movement | 30 cm (11.8 inch) | DHF357242 | US.Pos.Clea.MovementIDshift - Stroke | Maikel van Eekelen
Geometry | - | x | - | - | x | - | - | - | x | - | x | x | 1 | Frontal C-arm | Detector movement speed | 10 cm/sec (3.9 inch/sec) towards the patient, 15 cm/sec (5.9 inch/sec) away from the patient. | DHF357242 | US.Pos.Clea.MovementIDshift - Motorized Speed | Maikel van Eekelen
Geometry | - | x | - | - | x | - | - | - | x | - | x | x | 1 | Frontal C-arm | Detector Rotation | Floor stand: 0˚ to +90˚, rotation around the center of the functional detector surface Floor stand with L-arc-N 15 (Azurion R1.x, R2.x systems): 0˚ to +90˚, rotation around the center of the functional detector surface Floor stand with L-arc-N 15 (Azurion R3.0 system): -180˚ to +90˚, rotation around the center of the functional detector surface Floor stand with L-arc-C 12, L-arc-C 15, L-arc-N 12: 0˚ to +90˚, rotation around the center of the functional detector surface | DHF357242 | US.Pos.Clea.MovementFDRotation - Stroke | Maikel van Eekelen
Geometry | - | x | - | - | x | - | - | - | x | - | x | x | 1 | Frontal C-arm | Detector Rotation speed | Re-positioning of the Flat Detector from portrait to landscape (and vice versa) within 4 seconds | DHF357242 | US.Pos.Clea.MovementFDrotation - Motorized Speed | Maikel van Eekelen
Geometry | - | x | - | - | x | - | - | - | x | - | x | x | 1 | Frontal C-arm | Z-rotation | Floor stand: -90˚ to +90˚, Floor stand with L-arc-N 15 (Azurion R1.x, R2.x systems): -90˚ to +90˚ Floor stand with L-arc-N 15 (Azurion R3.0 system): -135˚ to +90˚ Floor stand with L-arc-C 12, L-arc-C 15, L-arc-N 12: -90˚ to +90˚ | DHF357242 | US.Pos.Clea.MovementZRotation - Stroke | Maikel van Eekelen
Geometry | - | x | - | - | x | - | - | - | x | - | x | x | 1 | Frontal C-arm | Z-rotation speed | 12°/sec. | DHF357242 | US.Pos.Clea.MovementZRotation - Motorized Speed | Maikel van Eekelen
Geometry | - | x | - | - | x | - | - | - | x | - | x | x | 1 | Frontal C-arm | In case of rotational scan | -185˚ to +120˚, maximum rotation speed in head position 55°/sec | DHF357242 | US.Pos.Clea.MovementRotationAngulation - High speed RA parameters | Maikel van Eekelen
Geometry | - | x | - | - | x | - | - | - | x | - | x | x | 1 | Frontal C-arm | Ceiling height | Floor stand: 230 cm (90.6 inch), Floor stand with LARC: 298 cm (117.3 inch) | DHF357242 | US.Pos.Clea.MainDimensionsClea - Clea floor US.Pos.Clea.MainDimensionsClea - Clea floor with LARC | Maikel van Eekelen
x |  | x |  |  | x |  |  |  | x |  | x | x
Geometry | - | - | - | - | - | - | - | - | - | x | x | x | 1 | Lateral L-arc | Lateral Stand for Biplane Systems | Tag: US.TechData.Lateral_Larc | - | - | -
Geometry | - | - | - | - | - | - | - | - | - | x | x | x | 1 | Lateral L-arc | Iso-center to floor | For B12/12: 106.5 cm (41.9 inch), For B20/12: 114 cm (44.9 inch), For B20/15: 114 cm (44.9 inch). | DHF207230 | US.POS.Larc.Functional.Dimensional | Maikel van Eekelen
Geometry | - | - | - | - | - | - | - | - | - | x | x | x | 1 | Lateral L-arc | L-arc rotation | L-arc-C with FD12: 0° LAO to 90° LAO, L-arc-C with FD15: 0° LAO to 90° LAO, L-arc-N with FD12: 27° RAO to 115° RAO, L-arc-N with FD15 (Azurion R1.x, R2.x systems): 27° RAO to 115° RAO, L-arc-N with FD15 (Azurion R3.0 system): 27° RAO to 117° RAO | DHF207230, DHF207230, DHF207230, DHF207230, DHF344806 | US.POS.LARC.Functional.RotateBeam - Stroke, US.POS.LARC.Functional.RotateBeam - Stroke, US.POS.LARC.Functional.RotateBeam - Stroke, US.POS.LARC.Functional.RotationAngulation - Stroke | Maikel van Eekelen
Geometry | - | - | - | - | - | - | - | - | - | x | x | x | 1 | Lateral L-arc | L-arc rotation speed | L-arc-C with FD12: 8°/sec, L-arc-C with FD15: 8°/sec, L-arc-N with FD12: 8°/sec, L-arc-N with FD15 (Azurion R1.x, R2.x systems): 8°/sec. L-arc-N with FD15 (Azurion R3.0 system): 16°/sec. | DHF207230, DHF207230, DHF207230, DHF207230, DHF344806 | US.POS.LARC.Functional.RotateBeam - Speed, US.POS.LARC.Functional.RotateBeam - Speed, US.POS.LARC.Functional.RotateBeam - Speed, US.POS.LARC.Functional.RotationAngulation - MotorizedSpeed | Maikel van Eekelen
Geometry | - | - | - | - | - | - | - | - | - | x | x | x | 1 | Lateral L-arc | L-arc angulation | 45° cranial to 45° caudal, possible at any rotation angle | DHF207230, DHF207230, DHF207230, DHF344806 | US.POS.LARC.Functional.AngulateBeam - Stroke, US.POS.LARC.Functional.AngulateBeam - Stroke, US.POS.LARC.Functional.AngulateBeam - Stroke, US.POS.LARC.Functional.RotationAngulation - Stroke | Maikel van Eekelen
Geometry | - | - | - | - | - | - | - | - | - | x | x | x | 1 | Lateral L-arc | L-arc angulation speed | L-arc-C with FD12: 8°/sec, L-arc-C with FD15: 8°/sec, L-arc-N with FD12: 8°/sec, L-arc-N with FD15 (Azurion R1.x, R2.x systems): 8°/sec. L-arc-N with FD15 (Azurion R3.0 system): 16°/sec. | DHF207230, DHF207230, DHF207230, DHF207230, DHF344806 | US.POS.LARC.Functional.AngulateBeam - Speed, US.POS.LARC.Functional.AngulateBeam - Speed, US.POS.LARC.Functional.AngulateBeam - Speed, US.POS.LARC.Functional.RotationAngulation - MotorizedSpeed | Maikel van Eekelen
Geometry | - | - | - | - | - | - | - | - | - | x | x | x | 1 | Lateral L-arc | Source-image distance | L-arc-C with FD12: 87.8 cm (34.6 inch) - 130.6 cm (51.4 inch), L-arc-C with FD15: 87.4 cm (34.4 inch) - 130.2 cm (51.3 inch), L-arc-N with FD12: 87.8 cm (34.6 inch) - 130.6 cm (51.4 inch), L-arc-N with FD15 (Azurion R1.x, R2.x systems): 87.4 cm (34.4 inch) - 130.2 cm (51.3 inch). L-arc-N with FD15 (Azurion R3.0 systems): 87.5 cm (34.4 inch) - 130.1 cm (51.2 inch). | DHF207230, DHF207230, DHF207230, DHF207230, DHF344806 | US.POS.LARC.Functional.IDShift - Stroke, US.POS.LARC.Functional.IDShift - Stroke, US.POS.LARC.Functional.IDShift - Stroke, US.POS.LARC.Functional.IDShift - Minimum SID, US.POS.LARC.Functional.IDShift - Stroke | Maikel van Eekelen
Geometry | - | - | - | - | - | - | - | - | - | x | x | x | 1 | Lateral L-arc | Longitudinal movement | L-arc-C with FD12: 315 cm (124 inch), optionally additional 150 cm (59.1 inch), L-arc-C with FD15: 315 cm (124 inch), optionally additional 150 cm (59.1 inch), L-arc-N with FD12: 315 cm (124 inch), optionally additional 150 cm (59.1 inch), L-arc-N with FD15 (Azurion R1.x, R2.x systems): 315 cm (124 inch), optionally additional 150 cm (59.1 inch), L-arc-N with FD15 (Azurion R3.0 systems): 312 cm (122.8 inch), optionally additional 146 cm (57.5 inch). | DHF207230, DHF207230, DHF207230, DHF207230, DHF344806 | US.POS.LARC.Functional.Longitudinal - Stroke (-extended), US.POS.LARC.Functional.Longitudinal - Stroke (-extended), US.POS.LARC.Functional.Longitudinal - Stroke (-extended), US.POS.LARC.Functional.Longitudinal - Stroke (-extended). | Maikel van Eekelen
Geometry | - | - | - | - | - | - | - | - | - | x | x | x | 1 | Lateral L-arc | Longitudinal movement speed | L-arc-C with FD12: 6 cm/sec (2.4 inch/sec) inside work area, 12 cm/sec (4.7 inch/sec) outside work area, L-arc-C with FD15: 6 cm/sec (2.4 inch/sec) inside work area, 12 cm/sec (4.7 inch/sec) outside work area, L-arc-N with FD12: 8 cm/sec (3.1 inch/sec) inside work area, 12 cm/sec (4.7 inch/sec) outside work area, L-arc-N with FD15 (Azurion R1.x, R2.x systems): 8 cm/sec (3.1 inch/sec) inside work area, 12 cm/sec (4.7 inch/sec) outside work area, L-arc-N with FD15 (Azurion R3.0 systems): 15 cm/sec (5.9 inch/sec), safe speed 4 cm/sec (1.6 inch/sec). | DHF207230, DHF207230, DHF207230, DHF207230, DHF344806 | US.POS.LARC.Functional.Longitudinal - Speed, US.POS.LARC.Functional.Longitudinal - Speed, US.POS.LARC.Functional.Longitudinal - Speed, US.POS.LARC.Functional.Longitudinal - Motorized speed | Maikel van Eekelen
Geometry | - | - | - | - | - | - | - | - | - | x | x | x | - | Lateral L-arc | Detector movement | L-arc-C with FD12: 42.8 cm (16.9 inch), L-arc-C with FD15: 42.8 cm (16.9 inch), L-arc-N with FD12: 42.8 cm (16.9 inch), L-arc-N with FD15 (Azurion R1.x, R2.x systems): 42.8 cm (16.9 inch), L-arc-N with FD15 (Azurion R3.0 systems): 42.6 cm (16.8 inch). | DHF207230, DHF207230, DHF207230, DHF207230, DHF344806 | US.POS.LARC.Functional.ID-Shift - Stroke , US.POS.LARC.Functional.ID-Shift - Stroke, US.POS.LARC.Functional.ID-Shift - Stroke, US.POS.LARC.Functional.FD-Shift - Stroke. | Maikel van Eekelen
Geometry | - | - | - | - | - | - | - | - | - | x | x | x | - | Lateral L-arc | Detector movement speed | L-arc-C with FD12: 6 cm/sec (2.4 inch/sec) towards the patient, 9 cm/sec (3.5 inch/sec) away from the patient, L-arc-C with FD15: 6 cm/sec (2.4 inch/sec) towards the patient, 9 cm/sec (3.5 inch/sec) away from the patient, L-arc-N with FD12: 6 cm/sec (2.4 inch/sec) towards the patient, 9 cm/sec (3.5 inch/sec) away from the patient, L-arc-N with FD15 (Azurion R1.x, R2.x systems): 6 cm/sec (2.4 inch/sec) towards the patient, 9 cm/sec (3.5 inch/sec) away from the patient, L-arc-N with FD15 (Azurion R3.0 systems): 10 cm/sec (3.9 inch/sec) towards the patient, 20 cm/sec (7.9 inch/sec) away from the patient. | DHF207230, DHF207230, DHF207230, DHF207230, DHF344806 | US.POS.LARC.Functional.ID-Shift - Speed, US.POS.LARC.Functional.ID-Shift - Speed, US.POS.LARC.Functional.ID-Shift - Speed, US.POS.LARC.Functional.FD-Shift - Motorized speed. | Maikel van Eekelen
x | x | [5 empty] | x |  |  | x
Geometry | x | - | - | - | - | - | x | - | - | x | - | - | 1 | Frontal C-arm | Floor-Mounted Stand with 12-Inch Detector | Tag: US.TechData.FrontalC-Arm.PolyG_Floor | - | - | -
Geometry | x | - | - | - | - | - | x | - | - | x | - | - | 1 | Frontal C-arm | Iso-center to floor | 106.5 cm (41.9 inch) | DHF329224 | US.Pos.polyG.Functional.Movements.Dimensions - Figure 4 | Maikel van Eekelen
Geometry | x | - | - | - | - | - | x | - | - | x | - | - | 1 | Frontal C-arm | C-arm rotation | 120° LAO, 120° RAO | DHF329224 | US.Pos.polyG.Functional.Movements.Propeller - Stroke | Maikel van Eekelen
Geometry | x | - | - | - | - | - | x | - | - | x | - | - | 1 | Frontal C-arm | C-arm rotation speed | 25°/s | DHF329224 | US.Pos.polyG.Functional.Movements.Propeller - Motorized speed | Maikel van Eekelen
Geometry | x | - | - | - | - | - | x | - | - | x | - | - | 1 | Frontal C-arm | In case of rotational angio | Maximum rotation speed 55°/sec. Rotation angle -120° to +120° | DHF329224 | US.Pos.polyG.Functional.Movements.Propeller - RA speed | Maikel van Eekelen
Geometry | x | - | - | - | - | - | x | - | - | x | - | - | 1 | Frontal C-arm | C-arm angulation | 45° cranial, 45° caudal | DHF329224 | US.Pos.polyG.Functional.Movements.Roll - Stroke Angle | Maikel van Eekelen
Geometry | x | - | - | - | - | - | x | - | - | x | - | - | 1 | Frontal C-arm | C-arm angulation speed | 25°/s | DHF329224 | US.Pos.polyG.Functional.Movements.Roll - Motorized speed | Maikel van Eekelen
Geometry | x | - | - | - | - | - | x | - | - | x | - | - | 1 | Frontal C-arm | L-arm rotation | -105° to +105° | DHF329224 | US.Pos.polyG.Functional.Movements.Floor.Park - Stroke | Maikel van Eekelen
Geometry | x | - | - | - | - | - | x | - | - | x | - | - | 1 | Frontal C-arm | L-arm rotation speed | 12°/s | DHF329224 | US.Pos.polyG.Functional.Movements.Floor.Park - Motorized speed | Maikel van Eekelen
Geometry | x | - | - | - | - | - | x | - | - | x | - | - | 1 | Frontal C-arm | Detector movement | 34.5 cm (13.6 inch) | DHF329224 | US.Pos.polyG.Functional.Movements.IDShift - Stroke | Maikel van Eekelen
Geometry | x | - | - | - | - | - | x | - | - | x | - | - | 1 | Frontal C-arm | Detector movement speed | 10 cm/sec (3.9 inch/sec) towards the patient, 20 cm/sec (7.9 inch/sec) away from the patient, | DHF329224 | US.Pos.polyG.Functional.Movements.IDShift - Motorized speed | Maikel van Eekelen
Geometry | x | - | - | - | - | - | x | - | - | x | - | - | 1 | Frontal C-arm | Focal spot to iso-center | 76.5 cm (30.1 inch) | DHF329224 | US.Pos.polyG.Functional.Movements.Dimensions - Figure 4 | Maikel van Eekelen
Geometry | x | - | - | - | - | - | x | - | - | x | - | - | 1 | Frontal C-arm | Source Image Distance | 89-123.5 cm (35.0-48.6 inch) | DHF329224 | US.Pos.polyG.Functional.Movements.IDShift - SID | Maikel van Eekelen
Geometry | x | - | - | - | - | - | x | - | - | x | - | - | 1 | Frontal C-arm | C-arm depth | 105 cm (41.3 inch) | DHF329224 | US.Pos.polyG.Functional.Movements.Dimensions - Figure 4 | Maikel van Eekelen
x |  |  |  | x |  |  |  | x
Geometry | - | - | - | x | - | - | - | x | - | - | - | - | 1 | Frontal C-arm | Ceiling-Mounted Stand with 20-Inch Detector | Tag: US.TechData.FrontalC-Arm.Clea_Ceiling | - | - | -
Geometry | - | - | - | x | - | - | - | x | - | - | - | - | 1 | Frontal C-arm | Iso-center to floor | 106.5 cm (41.9 inch) | DHF357242 | US.Pos.Clea.MainDimensionsClea - Figure 3a | Maikel van Eekelen
Geometry | - | - | - | x | - | - | - | x | - | - | - | - | 1 | Frontal C-arm | C-arm rotation | In head-end position: 120° LAO, 185° RAO, in side position: 90° LAO, 90° RAO | DHF357242 | US.Pos.Clea.MovementRotationAngulation - Stroke | Maikel van Eekelen
Geometry | - | - | - | x | - | - | - | x | - | - | - | - | 1 | Frontal C-arm | C-arm rotation speed | 25°/sec. | DHF357242 | US.Pos.Clea.MovementRotationAngulation - Motorized speed | Maikel van Eekelen
Geometry | - | - | - | x | - | - | - | x | - | - | - | - | 1 | Frontal C-arm | C-arm angulation | In head-end position: 90° cranial, 90° caudal In side position: 185° cranial, 120° caudal | DHF357242 | US.Pos.Clea.MovementRotationAngulation - Stroke | Maikel van Eekelen
Geometry | - | - | - | x | - | - | - | x | - | - | - | - | 1 | Frontal C-arm | C-arm angulation speed | 25°/sec. | DHF357242 | US.Pos.Clea.MovementRotationAngulation - Motorized speed | Maikel van Eekelen
Geometry | - | - | - | x | - | - | - | x | - | - | - | - | 1 | Frontal C-arm | Longitudinal movement | Ceiling Stand: 260 cm (102.4 inch) or 410 cm (161.4 inch) with an Extended rail of 150 cm (59.1 inch), FlexArm: 285 cm (112.2 inch), 460 cm (181.1 inch) with extended rail type 1 and 635 cm (250.0 inch) for extended rail type 2, FlexMove: 440 cm (173.2 inch) or 540 cm (212.6 inch) with extended rail option. | DHF357242 | US.Pos.Clea.MovementLongitudinalLateral - Stroke Y | Maikel van Eekelen
Geometry | - | - | - | x | - | - | - | x | - | - | - | - | 1 | Frontal C-arm | Lateral movement | FlexArm: 150 cm (59.1 inch), FlexMove: -130 cm (-51.2 inch) to +130 cm (+51.2 inch), relative to the centerline of the patient table. | DHF357242 | US.Pos.Clea.MovementLongitudinalLateral - Stroke X | Maikel van Eekelen
Geometry | - | - | - | x | - | - | - | x | - | - | - | - | 1 | Frontal C-arm | Longitudinal/Lateral movement speed | 15 cm/sec (5.9 inch/sec) | DHF357242 | US.Pos.Clea.MovementRotationAngulation - Motorized speed | Maikel van Eekelen
Geometry | - | - | - | x | - | - | - | x | - | - | - | - | 1 | Frontal C-arm | Focal spot to iso-center | 81 cm (31.9 inch) | DHF357242 | US.Pos.Clea.MainDimensionsClea - Figure 3a | Maikel van Eekelen
Geometry | - | - | - | x | - | - | - | x | - | - | - | - | 1 | Frontal C-arm | Source Image Distance | 89.5-119.5 cm (35.2-47.0 inch) | DHF357242 | US.Pos.Clea.MainDimensionsClea - Figure 3a | Maikel van Eekelen
Geometry | - | - | - | x | - | - | - | x | - | - | - | - | 1 | Frontal C-arm | C-arm depth | 90 cm (35.4 inch) | DHF357242 | US.Pos.Clea.MainDimensionsClea - Figure 3a | Maikel van Eekelen
Geometry | - | - | - | x | - | - | - | x | - | - | - | - | 1 | Frontal C-arm | Detector movement | 30 cm (11.8 inch) | DHF357242 | US.Pos.Clea.MovementIDshift - Stroke | Maikel van Eekelen
Geometry | - | - | - | x | - | - | - | x | - | - | - | - | 1 | Frontal C-arm | Detector movement speed | 10 cm/sec (3.9 inch/sec) towards the patient, 15 cm/sec (5.9 inch/sec) away from the patient. | DHF357242 | US.Pos.Clea.MovementIDshift - Motorized speed | Maikel van Eekelen
Geometry | - | - | - | x | - | - | - | x | - | - | - | - | 1 | Frontal C-arm | Detector Rotation | Ceiling Stand and Flexmove: 0˚ to +90˚, rotation around the center of the functional detector surface FlexArm: -180˚ to +90˚, rotation around the center of the functional detector surface | DHF357242 | US.Pos.Clea.MovementFDRotation - Stroke | Maikel van Eekelen
Geometry | - | - | - | x | - | - | - | x | - | - | - | - | 1 | Frontal C-arm | Detector Rotation speed | Re-positioning of the Flat Detector from portrait to landscape (and vice versa) within 4 seconds | DHF357242 | US.Pos.Clea.MovementFDrotation - Motorized Speed | Maikel van Eekelen
Geometry | - | - | - | x | - | - | - | x | - | - | - | - | 1 | Frontal C-arm | Z-rotation | Ceiling Stand: -90˚ to +90˚, FlexArm: -135˚ to +135˚, FlexMove: -90˚ to +90˚. | DHF357242 | US.Pos.Clea.MovementZRotation - Stroke | Maikel van Eekelen
Geometry | - | - | - | x | - | - | - | x | - | - | - | - | 1 | Frontal C-arm | Z-rotation speed | 12°/sec. | DHF357242 | US.Pos.Clea.MovementZRotation - Motorized Speed | Maikel van Eekelen
Geometry | - | - | - | x | - | - | - | x | - | - | - | - | 1 | Frontal C-arm | C-arm rotation in case of rotational scan | -185˚ to +120˚, maximum rotation speed in head position 55°/sec | DHF357242 | US.Pos.Clea.MovementRotationAngulation - High speed RA parameters | Maikel van Eekelen
Geometry | - | - | - | x | - | - | - | x | - | - | - | - | 1 | Frontal C-arm | Ceiling height | 270 cm (106 inch), with short L-arm, 290 cm (114 inch), with long L-arm 7C20 with FlexMove: 290 cm (114 inch), with short L-arm, 7C20 with FlexMove: 310 cm (122 inch), with long L-arm | DHF357242 | US.Pos.Clea.MainDimensionsClea - Clea floor US.Pos.Clea.MainDimensionsClea - Clea floor with LARC | Maikel van Eekelen
x
Geometry | - | - | x | - | - | x | - | - | - | - | - | - | 1 | Frontal C-arm | Ceiling-Mounted Stand with 12-Inch Detector | Tag: US.TechData.FrontalC-Arm.PolyG_Ceiling | - | - | -
Geometry | - | - | x | - | - | x | - | - | - | - | - | - | 1 | Frontal C-arm | Iso-center to floor | 106.5 cm (41.9 inch) | DHF329224 | US.Pos.polyG.Functional.Movements.Dimensions - Figure 4 | Maikel van Eekelen
Geometry | - | - | x | - | - | x | - | - | - | - | - | - | 1 | Frontal C-arm | Longitudinal movement | 260 cm (102.4 inch), or 410 cm (161.4 inch) with optional extended rail | DHF329224 | US.Pos.polyG.Functional.Movements.Longitudinal - Minimum stroke | Maikel van Eekelen
Geometry | - | - | x | - | - | x | - | - | - | - | - | - | 1 | Frontal C-arm | Longitudinal movement speed | 15 cm/sec (5.9 inch/sec) | DHF329224 | US.Pos.polyG.Functional.Movements.Longitudinal - Motorized speed | Maikel van Eekelen
Geometry | - | - | x | - | - | x | - | - | - | - | - | - | 1 | Frontal C-arm | L-arm rotation | -90˚ to +90˚ | DHF329224 | US.Pos.polyG.Functional.Movements.LArmRotation - Stroke | Maikel van Eekelen
Geometry | - | - | x | - | - | x | - | - | - | - | - | - | 1 | Frontal C-arm | L-arm rotation speed | 12°/sec. | DHF329224 | US.Pos.polyG.Functional.Movements.LArmRotation - Motorized speed | Maikel van Eekelen
Geometry | - | - | x | - | - | x | - | - | - | - | - | - | 1 | Frontal C-arm | C-arm rotation | 120° LAO, 120° RAO | DHF329224 | US.Pos.polyG.Functional.Movements.Propeller - Stroke | Maikel van Eekelen
Geometry | - | - | x | - | - | x | - | - | - | - | - | - | 1 | Frontal C-arm | C-arm rotation speed | 25°/s | DHF329224 | US.Pos.polyG.Functional.Movements.Propeller - Motorized speed | Maikel van Eekelen
Geometry | - | - | x | - | - | x | - | - | - | - | - | - | 1 | Frontal C-arm | In case of rotational angio: | Maximum rotation speed 55°/sec. Rotation angle -120° to +120° | DHF329224 | US.Pos.polyG.Functional.Movements.Propeller - RA speed | Maikel van Eekelen
Geometry | - | - | x | - | - | x | - | - | - | - | - | - | 1 | Frontal C-arm | C-arm angulation | 45° cranial, 45° caudal | DHF329224 | US.Pos.polyG.Functional.Movements.Roll - Stroke Angle | Maikel van Eekelen
Geometry | - | - | x | - | - | x | - | - | - | - | - | - | 1 | Frontal C-arm | C-arm angulation speed | 25°/s | DHF329224 | US.Pos.polyG.Functional.Movements.Roll - Motorized speed | Maikel van Eekelen
Geometry | - | - | x | - | - | x | - | - | - | - | - | - | 1 | Frontal C-arm | Detector movement | 34.5 cm (13.6 inch) | DHF329224 | US.Pos.polyG.Functional.Movements.IDShift - Stroke | Maikel van Eekelen
Geometry | - | - | x | - | - | x | - | - | - | - | - | - | 1 | Frontal C-arm | Detector movement speed | 10 cm/sec (3.9 inch/sec) towards the patient, 20 cm/sec (7.9 inch/sec) away from the patient, | DHF329224 | US.Pos.polyG.Functional.Movements.IDShift - Motorized speed | Maikel van Eekelen
Geometry | - | - | x | - | - | x | - | - | - | - | - | - | 1 | Frontal C-arm | Focal spot to iso-center | 76.5 cm (30.1 inch) | DHF329224 | US.Pos.polyG.Functional.Movements.Dimensions - Figure 4 | Maikel van Eekelen
Geometry | - | - | x | - | - | x | - | - | - | - | - | - | 1 | Frontal C-arm | Source Image Distance | 89-123.5 cm (35.0-48.6 inch) | DHF329224 | US.Pos.polyG.Functional.Movements.Dimensions - Figure 4 | Maikel van Eekelen
Geometry | - | - | x | - | - | x | - | - | - | - | - | - | 1 | Frontal C-arm | C-arm depth | 105 cm (41.3 inch) | DHF329224 | US.Pos.polyG.Functional.Movements.Dimensions - Figure 4 | Maikel van Eekelen
Geometry | - | - | x | - | - | x | - | - | - | - | - | - | 1 | Frontal C-arm | Ceiling height | 270 cm (106 inch), with short L-arm 290 cm (114 inch), with long L-arm | DHF329224 | US.Pos.polyG.Functional.Movements.Dimensions - Figure 4 | Maikel van Eekelen
x | x | x | x | x | x | x | x | x | x | x | x | x
Geometry | x | x | x | x | x | x | x | x | x | x | x | x | 1 | Patient Table | Patient Table | Tag: US.TechData.PatientTable | - | - | -
Geometry | x | x | x | x | x | x | x | x | x | x | x | x | 1 | Patient Table | Table Height | 74-102 cm (29.1 inch - 40.4 inch) for Patient Table without tilt/cradle option, 79-104 cm (31.1 inch - 40.9 inch) for Patient Table with tilt/cradle option | DHF318026 | URS.Pos.AD7X.Functional.Movements.Height | Stan Rutten
Geometry | x | x | x | x | x | x | x | x | x | x | x | x | 1 | Patient Table | Table top length (incl. OP Rail) | 319 cm (125.6 inch) | DHF318026 | URS.Pos.AD7X.Functional.Tabletops | Stan Rutten
Geometry | x | x | x | x | x | x | x | x | x | x | x | x | 1 | Patient Table | Table top width | 50 cm (19.7 inch) | DHF318026 | URS.Pos.AD7X.Functional.Tabletops | Stan Rutten
Geometry | x | x | x | x | x | x | x | x | x | x | x | x | 1 | Patient Table | Longitudinal float range | 120 cm (47.2 inch) | DHF318026 | URS.Pos.AD7X.Functional.Movements.Longitudinal | Stan Rutten
Geometry | x | x | x | x | x | x | x | x | x | x | x | x | 1 | Patient Table | Lateral float range | 36 cm (14.2 inch) | DHF318026 | URS.Pos.AD7X.Functional.Movements.Lateral | Stan Rutten
Geometry | x | x | x | x | x | x | x | x | x | x | x | x | 1 | Patient Table | Max. table load | 275 kg (606 lbs) + 500 N additional force max. tabletop extension in case of CPR | DHF318026 | URS.Pos.AD7X.Functional.General.Load | Stan Rutten
Geometry | x | x | x | x | x | x | x | x | x | x | x | x | 1 | Patient Table | Max. patient weight | 250 kg (551 lbs) | DHF318026 | URS.Pos.AD7X.Functional.General.Load | Stan Rutten
Geometry | x | x | x | x | x | x | x | x | x | x | x | x | 1 | Patient Table | Table up/down speed | 30 mm/s (1.2 inch/s) | DHF318026 | URS.Pos.AD7X.Functional.Movements.Height | Stan Rutten
Geometry | x | x | x | x | x | x | x | x | x | x | x | x | 1 | Patient Table options | Pivot | -90˚/+180˚ or -180˚/+90˚ | DHF318026 | URS.Pos.AD7X.Functional.Movements.Pivot | Stan Rutten
Geometry | x | x | x | x | x | x | x | x | x | x | x | x | 1 | Patient Table options | Swivel (includes pivot) | Extended longitudinal range: 78.2 cm (30.8 inch), Table height increase: +8 cm (3.2 inch), Pivot range: -90°/+180° only | DHF318026 | URS.Pos.AD7X.Functional.Movements.Swivel URS.Pos.AD7X.Functional.Movements.Pivot URS.Pos.AD7X.Functional.Movements.Height | Stan Rutten
Geometry | x | x | x | x | x | x | x | x | x | x | x | x | 1 | Patient Table options | Tilt and Cradle | Tilting range: ±16.5° iso-centric, Cradle tilting range: ±15°, Table height increase: Min height +5 cm (+1.97 inch) Max height +2 cm (0.78 inch) | DHF318026 | URS.Pos.AD7X.Functional.Movements.Tilt, URS.Pos.AD7X.Functional.Movements.Cradle, URS.Pos.AD7X.Functional.Movements.Height | Stan Rutten
Geometry | x | x | x | x | x | x | x | x | x | x | x | x | 1 | Patient Table options | Tilt | Tilting range: ±16.5° iso-centric, Table height increase: Min height +5 cm (+1.97 inch) Max height +2 cm (0.78 inch) | DHF318026 | URS.Pos.AD7X.Functional.Movements.Tilt, URS.Pos.AD7X.Functional.Movements.Height | Stan Rutten
x | x | x | x | x | x | x | x | x | x | x | x | x
Geometry | x | x | x | x | x | x | x | x | x | x | x | x | 1 | Philips MCS | Movement Range Rail system | Tag: US.TechData.MCS.RailSystem | - | - | Rene van Weert
Geometry | x | x | x | x | x | x | x | x | x | x | x | x | 1 | Philips MCS | Rotation range | 360° | DHF310430 | - | Rene van Weert
Geometry | x | x | x | x | x | x | x | x | x | x | x | x | 1 | Philips MCS | Transversal movement | Over a distance of 293 cm (115.4 inch) | DHF310430 | - | Rene van Weert
Geometry | x | x | x | x | x | x | x | x | x | x | x | x | 1 | Philips MCS | Longitudinal movement | Over a distance of 330 cm (129.9 inch) | DHF310430 | - | Rene van Weert
Geometry | x | x | x | x | x | x | x | x | x | x | x | x | 1 | Philips MCS | Height movement | Motorized 32 cm (12.6 inch) or 52 cm (20.5 inch), only for 2 fold MCS rail | DHF310421 | MCS.Rail.RoomLayout.VerticalRanges | Rene van Weert
Geometry | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Philips MCS | Movement Range 2 and 3 Fold light boom systems | Tag: US.TechData.MCS.LightBoomSystem | - | - | Rene van Weert
Geometry | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Philips MCS | Rotation range | 315° | DHF310429 | - | Rene van Weert
Geometry | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Philips MCS | Reach | Maximal range in horizontal plane 201 cm (79.1 inch) | DHF310429 | - | Rene van Weert
Geometry | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Philips MCS | Height movement | 106 cm (41.7 inch), can be adjusted to less range | DHF310429 | MCS.CeilingHeights | Rene van Weert
Geometry | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Philips MCS | Movement Range 1 and 4 Fold boom systems | Tag: US.TechData.BoomSystem | - | - | Rene van Weert
Geometry | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Philips MCS | Rotation range | 350° | DHF310430 | - | Rene van Weert
Geometry | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Philips MCS | Reach | Maximal range in horizontal plane 241 cm (94.9 inch) | DHF310430 | - | Rene van Weert
Geometry | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Philips MCS | Height movement | 82 cm (32.3 inch), can be adjusted to less range | DHF310421 | MCS.Spring.RoomLayout.ViewingHeights | Rene van Weert
Geometry | x | x | x | x | x | x | x | x | x | x | x | x | 1 | Philips MCS | Dimensions | Tag: US.TechData.MCS.Dimensions | - | - | Rene van Weert
Geometry | x | x | x | x | x | x | x | x | x | x | x | x | 1 | Philips MCS | 2 Fold MCS on Philips boom frame + monitor | Width 1500 mm (59.1 inch), height 600 mm (23.6 inch), depth 313 mm (12.3 inch) | - | - | Rene van Weert
Geometry | x | x | x | x | x | x | x | x | x | x | x | x | 1 | Philips MCS | 3 Fold MCS on Philips boom frame + monitor | Width 1027 mm (40.4 inch), height 925 mm (36.4 inch), depth 318 mm (12.5 inch) | - | - | Rene van Weert
Geometry | x | x | x | x | x | x | x | x | x | x | x | x | 1 | Philips MCS | 4 Fold MCS on Philips boom frame + monitor | Width 1500 mm (59.1 inch), height 1020 mm (40.2 inch), depth 330 mm (13 inch) | - | - | Rene van Weert
Geometry | x | x | x | x | x | x | x | x | x | x | x | x | 1 | Philips MCS | 6 Fold MCS on Philips boom frame + monitor | Width 1500 mm (59.1 inch), height 1020 mm (40.2 inch), depth 430 mm (16.9 inch) | - | - | Rene van Weert
Geometry | x | x | x | x | x | x | x | x | x | x | x | x | 1 | Philips MCS | 1 Fold FlexVision XL on Philips boom frame + monitor | Width 1600 mm (63 inch), height 1020 mm (40.2 inch), depth 335 mm (13.2 inch) | - | - | Rene van Weert
Geometry | x | x | x | x | x | x | x | x | x | x | x | x | 1 | Philips MCS | 3 Fold FlexVision XL on Philips boom frame + monitor | Width 1600 mm (63 inch), height 1020 mm (40.2 inch), depth 435 mm (17.1 inch) | - | - | Rene van Weert
Geometry | x | x | x | x | x | x | x | x | x | x | x | x | 1 | Philips MCS | Actuator | Tag: US.TechData.MCS.Actuator | - | - | Rene van Weert
Geometry | x | x | x | x | x | x | x | x | x | x | x | x | 1 | Philips MCS | Mains voltage | 230 V | DHF310430 | MCS.Rail.Monitors.Frame | Rene van Weert
Geometry | x | x | x | x | x | x | x | x | x | x | x | x | 1 | Philips MCS | Mains frequency | 50 / 60 Hz | DHF310430 | MCS.Rail.Monitors.Frame | Rene van Weert
Geometry | x | x | x | x | x | x | x | x | x | x | x | x | 1 | Philips MCS | Maximum power consumption | 500 W | DHF310430 | MCS.Rail.Monitors.Frame | Rene van Weert
Geometry | x | x | x | x | x | x | x | x | x | x | x | x | 1 | Philips MCS | Maximum speed | 12 to 24 mm/s (0.47 in/s to 0.95 in/s), depending of load and direction of movement. | DHF310430 | MCS.Rail.Monitors.Frame | Rene van Weert
x | x | x | x | x | x | x | x | x | x | x | x | x
Geometry | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Ceiling-Suspended Radiation Shield | Ceiling-Suspended Radiation Shield | Tag: US.TechData.CeilingRadiationShield | - | - | -
Geometry | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Ceiling-Suspended Radiation Shield | Suspension arm | Two-section suspension arm, 75 cm (29.5 inch) horizontal beam, 91 cm (35.8 inch) angulated part | 989600132422 1910__ECC_01: MW111776 | no tag | Sudeendra Chitta
Geometry | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Ceiling-Suspended Radiation Shield | Radiation shield | 50 x 40 cm (19.7 x 15.7 inch) with rounded corners. Attached to this shield a flexible apron 50 x 30 cm (19.7 x 11.8 inch) is mounted. | 989600132422 1910__ECC_01: MW111777 | no tag | Sudeendra Chitta
Geometry | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Ceiling-Suspended Radiation Shield | Material of the radiation shield | Transparant leaded acrylic plate, X-ray protection equivalent of 0.5 mm (0.02 inch) Pb. Also applicable for the apron. | 989600132422 1910__ECC_01: MW111778 | no tag | Sudeendra Chitta
Geometry | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Ceiling-Suspended Radiation Shield | Accessory Bracket: | Tag: US.TechData.CeilingRadiationShieldBracket | 989600132422 1910__ECC_01: MW111779 | no tag | Sudeendra Chitta
Geometry | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Ceiling-Suspended Radiation Shield | Pivot pin length | 32 mm (1.26 inch) | 989600132422 1910__ECC_01: MW111780 | no tag | Sudeendra Chitta
Geometry | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Ceiling-Suspended Radiation Shield | Maximum torgue | Total torgue generated by the multiplication of the weights by distance from the centre of gravity of the two components to the pivot pin may never exceed 225 Nm in all positions. | 989600132422 1910__ECC_01: MW111781 | no tag | Sudeendra Chitta
x
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Touch screen module | Touch Screen Module | Tag: US.TechData.TouchScreenModule | DHF267148 | Unit Specification TSM | Marc Arends
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Touch screen module | Hardware: |  |  |  | Marc Arends
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Touch screen module | Display | 12.1 inch, IPS, 1280x800 resolution, 350 cd/m² | DHF267148 DHF242464 | US.TSM_HW.Performance.DisplaySize;US.TSM_HW.Performance.DisplayResolution ;US.TSM_HW.Performance.DisplayBrightness 991932812031_Sheet_190 _TSM_1_1.pdf, tag 2.1.1.1 Display/Video | Marc Arends
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Touch screen module | Touch | Capacitive multi-touch (with improved touch settings for immunity to fluids) | DHF242464 | 991932812031_Sheet_190 _TSM_1_1.pdf; 5.2 Touch-Screen device | Marc Arends
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Touch screen module | Operating system | • Processor Intel Atom Bay-trail E3826 @1.46GHz • Non-volatile memory 32 Gb, flash-disk • Volatile memory 4 Gb DRAM | DHF267148 | US.TSM_HW.Performance.ProcessorSpeed US.TSM_HW.Performance.NonVolatileMemorySize US.TSM_HW.Performance.VolatileMemorySize | Marc Arends
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Touch screen module | Ingress protection marking (IPX) | IPX4 | DHF267148 | US.TSM_HW.DesignConstraints.General.IngressProtectionLevel | Marc Arends
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Touch screen module | Mechanical dimensions | 205 mm x 303 mm x 36 mm (H x W x D) | DHF242464 | 991932812031_Sheet_190 _TSM_1_1.pdf; 4.1.1 Dimensions | Marc Arends
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Touch screen module | Weight | ≤1.9 kg (net) | DHF242464 | 991932812031_Sheet_190 _TSM_1_1.pdf; 4.1.2 Weight | Marc Arends
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Touch screen module | Electromagnetic compliance | IEC 60601-1-2 Ed 4.0 2014-02 | DHF267148 | 6.2.1.2. EMC_DesignInput | Marc Arends
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Touch screen module | Certification | UL/CSA/IEC-60601-1 (patient environment) | DHF267148 | US.TSM_HW.Miscellaneous.Safety | Marc Arends
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Touch screen module | Maximum number in the exam room | 2 (table-side and foot-end) | DHF267148 | 2.4. Configurations | Marc Arends
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Touch screen module | Maximum number connected to the Azurion system | 3 | DHF267148 | 2.4. Configurations | Marc Arends
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Touch screen module | Protection | TSM protection bar | DHF289318 | US TSM Swingarm; US-TSMSwingarm.Geometry.Stability | Marc Arends
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Touch screen module | Software: |  |  |  | Lukasz Berek
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Touch screen module | Acquisition setting | present | DHF341847 | UID.Allura.UI.X-rayAcqApp.TSMLiveTab.AcquisitionTask | Lukasz Berek
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Touch screen module | MultiModality capabilities – control and command-compatible applications such as IntraSight from one touch screen module | present | DHF341844 | UID.Allura.UI.TSMFramework.TSMApplications, UID.Allura.UI.TSMFramework.GenericExternalEquipmentButtons | Lukasz Berek
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Touch screen module | Image processing | present | DHF341847 | UID.Allura.UI.X-rayAcqApp.TSMLiveTab.ProcessingTask | Lukasz Berek
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Touch screen module | Full system Automatic Position Control - an unlimited number of stand and table positions can be stored and recalled from touchscreen module. | present | DHF341847 | UID.Allura.UI.X-rayAcqApp.TSMLiveTab.GeometryTask | Lukasz Berek
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Touch screen module | Quantitative analysis (QA) | present | DHF341847 | UID.Allura.UI.X-rayAcqApp.TSM.Measurements | Lukasz Berek
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Touch screen module | Tableside control of multiple Philips platforms such as IntelliSpace Cardiovascular, Hemodynamics, CX50, and interventional tools | present | DHF341844 | UID.Allura.UI.TSMFramework.TSMApplications | Lukasz Berek
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Touch screen module | Table and geometry lock functions | present | DHF341844 | UID.Allura.UI.X-rayAcqApp.TSMLiveTab.AcquisitionTask | Lukasz Berek
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Touch screen module | Fluoroscopy buzzer reset | present | DHF341844 | UID.Allura.UI.TSMFramework.FluoroBuzzerNotification | Lukasz Berek
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Touch screen module | Fluoro store | present | DHF341844 | UID.Allura.UI.TSMFramework.General | Lukasz Berek
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Touch screen module | Cleaning mode | present | DHF341844 | UID.Allura.UI.TSMFramework.TSMApplications | Lukasz Berek
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Touch screen module | Control of monitors (Switchable Viewing or FlexVision) | present | DHF341844 | UID.Allura.UI.FlexVisionWorkspot.TSMUI, UID.Allura.UI.ExamRoomWorkspot.SwitchableMonitors.TSMUI | Lukasz Berek
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Touch screen module | X-ray enable/disable | present | DHF341844 | UID.Allura.UI.TSMFramework.General | Lukasz Berek
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Touch screen module | Stopwatch | present | DHF341844 | UID.Allura.UI.TSMFramework.General | Lukasz Berek
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Touch screen module | Mouse keyboard application to perform tasks | present | DHF341844 | UID.Allura.UI.TSMFramework.MouseAndOnScreenKeyboardApplication | Lukasz Berek
x
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Touch screen module Pro | Touch Screen Module Pro | Tag: US.TechData.TouchScreenModulePro
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Touch screen module Pro | Visualization of the acquired images | present | DHF341847 | UID.Allura.UI.X-rayAcqApp.TSM.ImageViewport | Lukasz Berek
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Touch screen module Pro | Draw markers on the images directly on the TSM screen | present | DHF341847 | UID.Allura.UI.X-rayAcqApp.TSM.Markers | Lukasz Berek
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Touch screen module Pro | Adjust the collimator by drag and drop on the clinical images | present | DHF341847 | UID.Allura.UI.X-rayAcqApp.TSMLiveTab.CollimationTask | Lukasz Berek
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Touch screen module Pro | Use the interventional room pointer to highlight details on the clinical images | present | DHF341847 | UID.Allura.UI.X-rayAcqApp.TSM.ImagePointer | Lukasz Berek
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Touch screen module Pro | Zoom and Pan clinical images by finger pinching | present | DHF341844 | UID.Allura.UI.TSMFramework.Gestures | Lukasz Berek
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Touch screen module Pro | Pictorial index with a diagonal pattern to differentiate the unsaved fluoro images | present | DHF341847 | UID.Allura.UI.X-rayAcqApp.TSMLiveTab.SeriesTask | Lukasz Berek
x
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Control Module | Control Module | Tag: US.TechData.ControlModule | DHF272045 |  | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Control Module | Ingress Protection Marking (IPX) | IPX4 | DHF272045 | US.TSO-MIG.IPX4 | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Control Module | Control module size | Biplane | DHF272045 |  | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Control Module | Geo module | 37cm x 18cm without belly bar | DHF272045 | no requirement in Unit Spec | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Control Module | Biplane imaging module | 23cm x 18cm without belly bar | DHF272045 | no requirement in Unit Spec | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Control Module | Weight Geo module: | 4 kg, Table Clamp & Belly bar: 1 kg. | DHF272045 | US.TSO-MIG. Weight | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Control Module | Biplane imaging module: | 1.8 kg | DHF272045 | no requirement in Unit Spec | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Control Module | Maximum number of control modules: |  | DHF272045 |  | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Control Module | A. Philips AD7 table | • Geo (NT or Tilt) • Imaging • Geo ORT (optional in ER on pedestal) • Geo (NT or Tilt or ORT) (optional in CR) | DHF272045 | - | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Control Module | B. Maquet table | • Imaging • Geo ORT | DHF272045 | - | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Control Module | Backlighting | Available | DHF272045 | US.TSO-MIG. Button icons; US.TSO-MIG. Housing icons | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Control Module | Cleanability | Simple to clean. The control modules are made with the minimum amount of gaps and grooves possible, the top layer is made of a molded silicone type material that integrated the buttons and joysticks ensuring maximum protection for liquid and dirt ingress. The joystick can be removed to improve cleanability. | DHF272045 | cm.pc.design-for cleanab.us.ngui.cleanability.dm US.TSO-MIG.IPX4 US.TSO-MIG. Design US.TSO-MIG. Surface affection | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Control Module | Protective bar | Equipped with a belly bar for protection | DHF272045 | US.TSO-MIG. belly bar | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | - | Control Module | Cable length and cable extender option | A standard cable of 2.2m with additional cabling options of 8.0 m for ORT/ pedestal | DHF272045 | US.TSO-MIG. Standard Cable length: Cable length= 2200mm +/- 50mm excluding snap-fit relief length. US.TSO-MIG. ORT Cable length: The ORT TSO variants shall have a cable length of 8000mm +/- 50mm excluding snap-fit relief length | Kateryna Fisher Kulyk
x
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Viewpad | Viewpad |  | DHF215916 |  | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Viewpad | Viewpad Laser pointer | Tag: US.TechData.LaserData | DHF215916 |  | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Viewpad | Wavelength | Between 630 and 640 nm | DHF215916 | US.viewpad.laserpointer.wavelength | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Viewpad | Maximum output power | < 1mW | DHF215916 | US.viewpad.laserpointer.outputpower | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Viewpad | Viewpad Cardio functions | Tag: US.TechData.ViewpadFunctions | DHF215916 | US.viewpad.key definitions cardio | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Viewpad | 1 Copies the current image to the Reference 1 window. On biplane systems, the frontal image is copied. | present | DHF215916 | US.viewpad.key definitions cardio | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Viewpad | 2 Copies the current image to the Reference 2 window. On biplane systems, the lateral image is copied. | present | DHF215916 | US.viewpad.key definitions cardio | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Viewpad | 3 Copies the current image to the Reference 3 window. On biplane systems, images from both channels are copied, and are displayed side by side and synchronized. | present | DHF215916 | US.viewpad.key definitions cardio | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Viewpad | 4 Displays the previous series | present | DHF215916 | US.viewpad.key definitions cardio | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Viewpad | 5 Displays the next series | present | DHF215916 | US.viewpad.key definitions cardio | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Viewpad | 6 Displays the previous image | present | DHF215916 | US.viewpad.key definitions cardio | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Viewpad | 7 Displays the next image | present | DHF215916 | US.viewpad.key definitions cardio | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Viewpad | 8 Displays all series in the study overview | present | DHF215916 | US.viewpad.key definitions cardio | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Viewpad | 9 Plays the current series in looped movie mode | present | DHF215916 | US.viewpad.key definitions cardio | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Viewpad | 10 Plays all series of the study in looped movie mode | present | DHF215916 | US.viewpad.key definitions cardio | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Viewpad | 11 Moves the focus of the viewpad between the live X-ray window and each of the reference windows | present | DHF215916 | US.viewpad.key definitions cardio | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Viewpad | 12 Creates a snapshot of the current image and stores it with the study | present | DHF215916 | US.viewpad.key definitions cardio | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Viewpad | Viewpad Vascular functions | Tag: US.TechData.ViewpadFunctions | DHF215916 | US.viewpad.key definitions vascular | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Viewpad | 1 Copies the current image to the Reference 1 window. On biplane systems, the frontal image is copied. | present | DHF215916 | US.viewpad.key definitions vascular | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Viewpad | 2 Copies the current image to the Reference 2 window. On biplane systems, the lateral image is copied. | present | DHF215916 | US.viewpad.key definitions vascular | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Viewpad | 3 Copies the current image to the Reference 3 window. On biplane systems, images from both channels are copied, and are displayed side by side and synchronized. | present | DHF215916 | US.viewpad.key definitions vascular | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Viewpad | 4 Displays the previous series | present | DHF215916 | US.viewpad.key definitions vascular | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Viewpad | 5 Displays the next series | present | DHF215916 | US.viewpad.key definitions vascular | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Viewpad | 6 Displays the previous image | present | DHF215916 | US.viewpad.key definitions vascular | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Viewpad | 7 Displays the next image | present | DHF215916 | US.viewpad.key definitions vascular | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Viewpad | 8 Displays all series in the study overview | present | DHF215916 | US.viewpad.key definitions vascular | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Viewpad | 9 Plays the current series in looped movie mode | present | DHF215916 | US.viewpad.key definitions vascular | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Viewpad | 10 Plays all series of the study in looped movie mode | present | DHF215916 | US.viewpad.key definitions vascular | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Viewpad | 11 Moves the focus of the viewpad between the live X-ray window and each of the reference windows | present | DHF215916 | US.viewpad.key definitions vascular | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Viewpad | 12 Creates a snapshot of the current image and stores it with the study | present | DHF215916 | US.viewpad.key definitions vascular | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Viewpad | 13 Enables or disables subtraction | present | DHF215916 | US.viewpad.key definitions vascular | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Viewpad | 14 Sets the current image as the mask image for subtraction | present | DHF215916 | US.viewpad.key definitions vascular | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Viewpad | 15 Enables or disables landmarking | present | DHF215916 | US.viewpad.key definitions vascular | Kateryna Fisher Kulyk
x
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Wireless Foot Switch | Wireless Foot Switch | Tag: US.TechData.WFS | - | - | -
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Wireless Foot Switch | Frequency range | 2400 MHz to 2483.5 MHz | DHF700030 | - | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Wireless Foot Switch | Channel spacing | 2 MHz | DHF700030 | - | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Wireless Foot Switch | Modulation | Gaussian Frequency Shift Keying (GFSK) Adaptive frequency-hopping on 40 channels | DHF700030 | - | Kateryna Fisher Kulyk
User Interface | - | - | - | - | - | - | - | - | - | - | - | - | 1 | Wireless Foot Switch | Conformity | Europe: EN 300 328 v2.2.2, USA: FCC Part 15C, single modular, FCC license grant number: XK5 -SW24LE Canada: RSS247 Issue 1, IC license grant number: 5158A-SW24LE | DHF700030 | - | Kateryna Fisher Kulyk
x
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 3.1 X-ray generator | Technical Data Certeray X-ray Generator | Tag: US.TechData.XRayGenerator | - | - | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | - | 3.1 X-ray generator | Generated power | Max power 100 kW | GTC-DAT-0588-12 | - | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | - | 3.1 X-ray generator | Minimum switching time | Quartz-controlled power switch, with a minimum switching time of 1 ms | - | - | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | - | 3.1 X-ray generator | Voltage range | 40 to 125 kV | GTC-DAT-0588-12 | par 6.4, see Tube Voltage | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | - | 3.1 X-ray generator | Maximum current | 1000 mA at 100 kV | GTC-DAT-0588-12 | par 6.3 see Max electrical power | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | - | 3.1 X-ray generator | Maximum continuous power | 2.5 kW for 0.25 hours, 1.5 kW for 8 hours | - | - | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | - | 3.1 X-ray generator | Nominal power (highest electrical power) | 100 kW (1000 mA at 100 kV) | GTC-DAT-0588-12 | par 6.3 see Nominal electrical power | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 3.1 X-ray generator | Methods of Measurement |  | GTC-DAT-0588-12 | par 6.1 | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 3.1 X-ray generator | X-ray tube voltage | Tube voltage is measured with the aid of balanced high-voltage bleeders in the high voltage circuit, Measurement unit: Kilovolts (kV) | GTC-DAT-0588-12 | par 6.1 | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 3.1 X-ray generator | X-ray tube current | Tube current is measured on the cathode side in the rectified high-voltage circuit of the X-ray generator, Measurement unit: Milliamperes (mA) | GTC-DAT-0588-12 | par 6.1 | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 3.1 X-ray generator | Load time | Load time is measured between 75% ±7.5% peak voltage of the highvoltage rise edge and 75% ±7.5% peak voltage of the high-voltage fall edge, Measurement unit: Seconds (s) or milliseconds (ms) | GTC-DAT-0588-12 | par 6.1 | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 3.1 X-ray generator | Current-time product | Current-time product is measured on the cathode side in the rectified high voltage circuit of the HV generator between 75% ±7.5% peak voltage of the high-voltage rise edge and 75% ±7.5% peak voltage of the high-voltage fall edge, Measurement unit: Milliampereseconds (mAs) | GTC-DAT-0588-12 | par 6.1 | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 3.1 X-ray generator | Loading Factors |  | GTC-DAT-0588-12 | par 6.2 | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 3.1 X-ray generator | Maximum X-ray tube voltage and highest X-ray tube current at that voltage, Mode: Radiographic (Intermittent) | 125 kV, 720 mA | GTC-DAT-0588-12 | par 6.2 | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 3.1 X-ray generator | Maximum X-ray tube current and highest X-ray tube voltage at that current, Mode: Radiographic (Intermittent) | 1000 mA, 100 kV | GTC-DAT-0588-12 | par 6.2 | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 3.1 X-ray generator | Combination of X-ray tube current and X-ray tube voltage resulting in highest output power, Mode: Radiographic (Intermittent) | 1000 mA, 100 kV | GTC-DAT-0588-12 | par 6.2 | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 3.1 X-ray generator | Highest constant output power at 100 kV, 0.1s, Mode: Radiographic (Intermittent) | 100 kW, 1000 mA | GTC-DAT-0588-12 | par 6.2 | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 3.1 X-ray generator | The lowest current time product or the combination of loading factors resulting in the lowest current time product, Mode: Radiographic (Intermittent) | 0.1 mAs | GTC-DAT-0588-12 | par 6.2 | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 3.1 X-ray generator | Nominal shortest irradiation time (AEC exposures), Mode: AEC | Not applicable: no photo time technique | GTC-DAT-0588-12 | par 6.2 | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 3.1 X-ray generator | Ranges of tube load factors controlled by AEC, Mode: AEC | Ranges of tube load factors determined by X-ray protocol. The maximum range is: 40 – 125 kV; 10 – 1000 mA. Tolerances according to §203.6.4.3.104.3 and §203.6.4.3.104.4 | GTC-DAT-0588-12 | par 6.2 | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 3.1 X-ray generator | Electrical Data - Generator |  | GTC-DAT-0588-12 | par 6.3 | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 3.1 X-ray generator | Power supply For current catalog numbers: 722229, 722230, 722231, 722232, 722233, 722234, 722235, 722236, and 722220, For previous catalog numbers: 722221, 722222, 722223, 722224, 722225, 722226, 722227, and 722228 | 380 V – 480 V ±10%, 50 Hz / 60 Hz, 3 – phase, switched and fused (50 A slow blow) by system PDU | GTC-DAT-0588-12 | par 6.3 | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 3.1 X-ray generator | Power supply For previous catalog numbers: 722063, 722064, 722067, 722068, 722078, and 722079 | 400 V – 480 V ±10%, 50 Hz / 60 Hz, 3 – phase, switched and fused (50 A slow blow) by system PDU | GTC-DAT-0588-12 | par 6.3 | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 3.1 X-ray generator | Radiography Maximum voltage power | 125 kV | GTC-DAT-0588-12 | par 6.3 | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 3.1 X-ray generator | Radiography X-ray tube current | 1000 mA | GTC-DAT-0588-12 | par 6.3 | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 3.1 X-ray generator | Radiography Nominal electrical power | 100 kW (100 kV, 1000 mA, 0.1 s) | GTC-DAT-0588-12 | par 6.3 | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 3.1 X-ray generator | Radiography Maximum electrical powe | • 100 kW • 1000 mA at 100 kV | GTC-DAT-0588-12 | par 6.3 | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 3.1 X-ray generator | Continuous output | 1.5 kW (for example, 9 fpm at 100 kW, 0.1 s) | GTC-DAT-0588-12 | par 6.3 | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 3.1 X-ray generator | High-voltage generation | Converter | GTC-DAT-0588-12 | par 6.3 | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 3.1 X-ray generator | Ripple | DC voltage | GTC-DAT-0588-12 | par 6.3 | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 3.1 X-ray generator | Duty cycle | The generator can be used continuously as long as average power limitations described in X-ray System Configuration (page 316) are satisfied. | GTC-DAT-0588-12 | par 6.3 | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 3.1 X-ray generator | Radiography with Automatic Exposure Control |  | GTC-DAT-0588-12 | par 6.4 | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 3.1 X-ray generator | mAs | 0.01 mAs…10 mAs | GTC-DAT-0588-12 | par 6.4 | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 3.1 X-ray generator | Switching times | 3.0 ms…10 ms | GTC-DAT-0588-12 | par 6.4 | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 3.1 X-ray generator | Radiography without Automatic Exposure Control |  | GTC-DAT-0588-12 | par 6.4 | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 3.1 X-ray generator | Tube voltage | 40 kV…125 kV adjustable in steps of 1 kV or according to a sequence the steps of which roughly correspond to an exposure increment. In the case of tubes with lower maximum voltage this is limited accordingly. | GTC-DAT-0588-12 | par 6.4 | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 3.1 X-ray generator | Tube current | 10 m A...1000 mA, For kV-mA-s and kV-mAs techniques this can be adjusted in steps 1 | GTC-DAT-0588-12 | par 6.4 | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 3.1 X-ray generator | mAs | range 0.1 mAs…2000 mAs, Adjustable in steps 1 | GTC-DAT-0588-12 | par 6.4 | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 3.1 X-ray generator | Exposure times | 1 ms…16 s, Adjustable in steps 1 | GTC-DAT-0588-12 | par 6.4 | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 3.1 X-ray generator | Note 1: Steps selectable on system level |  | GTC-DAT-0588-12 | par 6.4 | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 3.1 X-ray generator | Pulsed fluoroscopy with Grid Control |  | GTC-DAT-0588-12 | par 6.4 | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 3.1 X-ray generator | Tube voltage | 40 kV - 125 kV | GTC-DAT-0588-12 | par 6.4 | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 3.1 X-ray generator | Tube current | 10 mA - 200 mA (depending on the tube configuration) | GTC-DAT-0588-12 | par 6.4 | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 3.1 X-ray generator | Accuracy |  | GTC-DAT-0588-12 | par 6.5 | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 3.1 X-ray generator | Tube voltage (requested by standard: ± 8%) | ± (5%) | GTC-DAT-0588-12 | par 6.5 | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 3.1 X-ray generator | Tube current-time product (requested by standard: ± (10% + 0.2 mAs)) | ± (3% + 0.2 mAs) | GTC-DAT-0588-12 | par 6.5 | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 3.1 X-ray generator | Tube current (requested by standard: ± 20%) | ± (5% + 1.0 mA) (Tp ≥ 35 ms) ± (8% + 1.0 mA) (1 < Tp < 35 ms) | GTC-DAT-0588-12 | par 6.5 | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 3.1 X-ray generator | Exposure time (requested by standard: ± (10% + 1 ms) | ± (10% + 1 ms) | GTC-DAT-0588-12 | par 6.5 | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 3.1 X-ray generator | mAs post-exposure display | ± (3% + 0.2 mAs) | GTC-DAT-0588-12 | par 6.5 | Jos van Tongeren
X-ray generation | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 3.1 X-ray generator | Post-exposure time display | ± (2% + 0.1 ms) | GTC-DAT-0588-12 | par 6.5 | Jos van Tongeren
x |  | x |  | x | x |  |  | x | x |  | x | x
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Assembly of X-ray tube and collimator (X-ray source assembly) | Tag: US.TechData.X-ray_Source_Assembly | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Loading factors corresponding to the maximum specified energy input to anode in one hour when applied at the nominal X-ray tube voltage | 125 kV, 28 mA (3500 W) | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Maximum symmetrical radiation field | MRC 200+ 0508 ROT-GS 1003: 28 x 28 cm at 1 m distance, MRC 200+ 0407 ROT-GS 1004: 35 x 35 cm at 1 m distance, MRC 200+ 0407 ROT-GS 1008: 35 x 35 cm at 1 m distance, MRC 200 0508 ROT-GS 1003: 28 x 28 cm at 1 m distance, MRC 200 0407 ROT-GS 1004: 35 x 35 cm at 1 m distance | - | - | Jonne Moeskops
x |  | x |  | x | x |  |  | x | x |  | x | x
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Technical Data MRC 200+ GS 0407 Tube | Tag: US.TechData.X-rayTube_MRC 200+_GS_0407 | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Focal spot size and loadability | 0.4 mm nominal focal spot with maximal 30 kW loadability 0.7 mm nominal focal spot with maximal 65 kW loadability Based on 250 W anode reference power | GTC-DAT-0163-15 | TRS001; TRS004 | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Grid-switched pulsed fluoroscopy | yes | GTC-DAT-0163-15 | TRS012; TRS027 | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Fluoro power for 10 minutes | 4,500 W | GTC-DAT-0163-15 | TRS042 (par 2.2) | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Fluoro power for 20 minutes | 4,000 W | GTC-DAT-0163-15 | TRS043 (par 2.2) | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Maximum X-ray field with SID = 100 | 35 x 35 cm (13.8 x 13.8 inch) | GTC-DAT-0163-15 | TRS015 (par 3.7) | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Maximum X-ray field with SID = 120 | 42 x 42 cm (16.5 x 16.5 inch) | GTC-DAT-0163-15 | TRS015 calculated | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Maximum X-ray field with SID = 70 | 24.5 x 24.5 cm (9.6 x 9.6 inch) | GTC-DAT-0163-15 | TRS015 calculated | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Maximum anode cooling rate | 1750 kHU/min | GTC-DAT-0163-15 | par 2.2 | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Max. anode heat storage | 6.4 MHUeff | GTC-DAT-0163-15 | par 2.2 | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Max. assembly heat storage | 9.4 MHUeff | GTC-DAT-0163-15 | par 2.2 | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Anode heat dissipation | 21,000 W | GTC-DAT-0163-15 | par 2.2 | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Continuous anode heat dissipation | 3,500 W | GTC-DAT-0163-15 | par 2.2 | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Max. assembly continuous heat dissipation | 4,000 W | GTC-DAT-0163-15 | par 2.2 | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | - | 3.2 X-ray tube | Extra pre-filtration | SpectraBeam filters: 0.1, 0.4, 0.9 mm Cu and 1 mm Al backing. | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | - | 3.2 X-ray tube | Cooling liquid | Oil cooled X-ray tube with thermal safety switch | GTC-DAT-0163-15 | TRS026 | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Anode target angle | 11° | GTC-DAT-0163-15 | TRS014 | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | - | 3.2 X-ray tube | Anode cooling method | Direct anode oil cooling system with 200 mm anode diameter | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Tube Output Power |  | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Exposure |  | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Highest electric power (X-ray tube current and X-ray tube voltage resulting in highest electric power) | 65 kW (125 kV, 520 mA) | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Highest electric output power at 100kV, 0.1s (X-ray tube current) | 65 kW (650 mA) | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Maximum X-ray tube voltage and highest X-ray tube current at that voltage | 125 kV, 520 mA | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Maximum X-ray tube current and highest X-ray tube voltage at that current | 813 mA, 80 kV | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Minimum X-ray tube current and lowest X-ray tube voltage at that current | 10 mA, 40 kV | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | The lowest current time product (loading factors at the lowest current time product) | 0.1 mAs (10 mA, 10 ms) | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Fluoroscopy with Grid Switch |  | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Highest electric power (X-ray tube current and X-ray tube voltage resulting in highest electric power) | 20 kW (125 kV, 160 mA) | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Highest electric output power at 100kV, 0.1s (X-ray tube current) | 16 kW (160 mA) | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Maximum X-ray tube voltage and highest X-ray tube current at that voltage | 125 kV1, 160 mA | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Minimum X-ray tube voltage and lowest X-ray tube current at that voltage | 40 kV, 10 mA | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Maximum X-ray tube current and highest X-ray tube voltage at that current | 160 mA, 125 kV | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Minimum X-ray tube current and lowest X-ray tube voltage at that current | 10 mA, 40 kV | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | The lowest current time product (loading factors at the lowest current time product) | 0.007 mAs (2 mA, 3.5 ms) | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Note 1: Normal usage is 120 kV. For low load fluoroscopy, 125 kV is possible. |  | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | General note: Fluoroscopy is available in pulsed fluoroscopy mode only. |  | - | - | Jonne Moeskops
x |  | x |  | x | x |  |  | x | x |  | x | x
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Technical Data MRC 200+ GS 0407 Tube | Tag: US.TechData.X-rayTube_MRC 200+_GS_0407_1008 | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Focal spot size and loadability | 0.4 mm nominal focal spot with maximal 30 kW loadability 0.7 mm nominal focal spot with maximal 65 kW loadability Based on 250 W anode reference power | GTC-DAT-0028-15 | TRS1103; TRS1104 | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Grid-switched pulsed fluoroscopy | yes | GTC-DAT-0028-15 | TRS1113; TRS1114 | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Fluoro power for 10 minutes | 4,500 W | GTC-DAT-0028-15 | TRS1319 (par 2.2) | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Fluoro power for 20 minutes | 4,000 W | GTC-DAT-0028-15 | TRS1320 (par 2.2) | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Maximum X-ray field with SID = 100 | 35 x 35 cm (13.8 x 13.8 inch) | GTC-DAT-0028-15 | TRS1614 (par 3.7) | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Maximum X-ray field with SID = 120 | 42 x 42 cm (16.5 x 16.5 inch) | GTC-DAT-0028-15 | TRS1614 calculated | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Maximum X-ray field with SID = 70 | 24.5 x 24.5 cm (9.6 x 9.6 inch) | GTC-DAT-0028-15 | TRS1614 calculated | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Maximum anode cooling rate | 1750 kHU/min | GTC-DAT-0028-15 | TRS1124 (par 2.2) | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Max. anode heat storage | 6.4 MHUeff | GTC-DAT-0028-15 | TRS1122 (par 2.2) | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Max. assembly heat storage | 9.4 MHUeff | GTC-DAT-0028-15 | TRS1313 (par 2.2) | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Anode heat dissipation | 21,000 W | GTC-DAT-0028-15 | TRS1124 (par 2.2) | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Continuous anode heat dissipation | 3,500 W | GTC-DAT-0028-15 | TRS1123 (par 2.2) | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Max. assembly continuous heat dissipation | 4,000 W | GTC-DAT-0028-15 | TRS1318 (par 2.2) | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | - | 3.2 X-ray tube | Extra pre-filtration | SpectraBeam filters: 0.1, 0.4, 0.9 mm Cu and 1 mm Al backing. | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | - | 3.2 X-ray tube | Cooling liquid | Oil cooled X-ray tube with thermal safety switch | GTC-DAT-0028-15 | TRS1841 | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Anode target angle | 11° | GTC-DAT-0028-15 | TRS1121 | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | - | 3.2 X-ray tube | Anode cooling method | Direct anode oil cooling system with 200 mm anode diameter | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Tube Output Power |  | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Exposure |  | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Highest electric power (X-ray tube current and X-ray tube voltage resulting in highest electric power) | 65 kW (125 kV, 520 mA) | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Highest electric output power at 100kV, 0.1s (X-ray tube current) | 65 kW (650 mA) | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Maximum X-ray tube voltage and highest X-ray tube current at that voltage | 125 kV, 520 mA | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Maximum X-ray tube current and highest X-ray tube voltage at that current | 813 mA, 80 kV | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Minimum X-ray tube current and lowest X-ray tube voltage at that current | 10 mA, 40 kV | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | The lowest current time product (loading factors at the lowest current time product) | 0.1 mAs (10 mA, 10 ms) | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Fluoroscopy with Grid Switch |  | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Highest electric power (X-ray tube current and X-ray tube voltage resulting in highest electric power) | 20 kW (125 kV, 160 mA) | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Highest electric output power at 100kV, 0.1s (X-ray tube current) | 16 kW (160 mA) | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Maximum X-ray tube voltage and highest X-ray tube current at that voltage | 125 kV1, 160 mA | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Minimum X-ray tube voltage and lowest X-ray tube current at that voltage | 40 kV, 10 mA | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Maximum X-ray tube current and highest X-ray tube voltage at that current | 160 mA, 125 kV | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Minimum X-ray tube current and lowest X-ray tube voltage at that current | 10 mA, 40 kV | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | The lowest current time product (loading factors at the lowest current time product) | 0.007 mAs (2 mA, 3.5 ms) | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Note 1: Normal usage is 120 kV. For low load fluoroscopy, 125 kV is possible. |  | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | General note: Fluoroscopy is available in pulsed fluoroscopy mode only. |  | - | - | Jonne Moeskops
x | x |  | x |  |  | x | x |  |  |  | x | x
X-ray generation | x | - | x | - | - | x | x | - | - | x | x | x | 1 | 3.2 X-ray tube | Technical Data MRC 200+ GS 0508 Tube | Tag: US.TechData.X-rayTube_MRC 200+_GS_0508 | - | - | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Focal spot size and loadability | 0.5 mm nominal focal spot with maximal 45 kW loadability 0.8 mm nominal focal spot with maximal 85 kW loadability Based on 250 W anode reference power | GTC-DAT-0260-15 | TRS001; TRS004 | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Grid-switched pulsed fluoroscopy | yes | GTC-DAT-0260-15 | TRS012; TRS027 | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Fluoro power for 10 minutes | 4,500 W | GTC-DAT-0260-15 | TRS042 (par 2.2) | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Fluoro power for 20 minutes | 4,000 W | GTC-DAT-0260-15 | TRS043 (par 2.2) | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Maximum X-ray field with SID = 100 | 28 x 28 cm (11.0 x 11.0 inch) | GTC-DAT-0260-15 | TRS015 (par 3.7) | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Maximum X-ray field with SID = 120 | 33,6 x 33,6 cm (13.2 x 13.2 inch) | GTC-DAT-0260-15 | TRS015 calculated | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Maximum X-ray field with SID = 70 | 19,6 x 19,6 cm (7.7 x 7.7 inch) | GTC-DAT-0260-15 | TRS015 calculated | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Maximum anode cooling rate | 1750 kHU/min | GTC-DAT-0260-15 | par 2.2 | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Max. anode heat storage | 6.4 MHUeff | GTC-DAT-0260-15 | par 2.2 | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Max. assembly heat storage | 9.4 MHUeff | GTC-DAT-0260-15 | par 2.2 | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Anode heat dissipation | 21,000 W | GTC-DAT-0260-15 | par 2.2 | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Continuous anode heat dissipation | 3,500 W | GTC-DAT-0260-15 | par 2.2 | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Max. assembly continuous heat dissipation | 4,000 W | GTC-DAT-0260-15 | par 2.2 | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | - | 3.2 X-ray tube | Extra pre-filtration | SpectraBeam filters: 0.1, 0.4, 0.9 mm Cu and 1 mm Al backing. | - | - | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | - | 3.2 X-ray tube | Cooling liquid | Oil cooled X-ray tube with thermal safety switch | GTC-DAT-0260-15 | TRS026 | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Anode target angle | 9° | GTC-DAT-0260-15 | TRS014 | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | - | 3.2 X-ray tube | Anode cooling method | Direct anode oil cooling system with 200 mm anode diameter | - | - | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Tube Output Power |  | - | - | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Exposure |  | - | - | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Highest electric power (X-ray tube current and X-ray tube voltage resulting in highest electric power) | 85 kW (125 kV, 680 mA) | - | - | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Highest electric output power at 100kV, 0.1s (X-ray tube current) | 85 kW (850 mA) | - | - | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Maximum X-ray tube voltage and highest X-ray tube current at that voltage | 125 kV, 680 mA | - | - | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Maximum X-ray tube current and highest X-ray tube voltage at that current | 1063 mA, 80 kV (The maximum tube current cannot be reached with the current system configuration) | - | - | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Minimum X-ray tube current and lowest X-ray tube voltage at that current | 10 mA, 40 kV | - | - | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | The lowest current time product (loading factors at the lowest current time product) | 0.1 mAs (10 mA, 10 ms) | - | - | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Fluoroscopy with Grid Switch |  | - | - | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Highest electric power (X-ray tube current and X-ray tube voltage resulting in highest electric power) | 25 kW (125 kV, 200 mA) | - | - | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Highest electric output power at 100kV, 0.1s (X-ray tube current) | 20 kW (200 mA) | - | - | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Maximum X-ray tube voltage and highest X-ray tube current at that voltage | 125 kV1, 200 mA | - | - | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Minimum X-ray tube voltage and lowest X-ray tube current at that voltage | 40 kV, 10 mA | - | - | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Maximum X-ray tube current and highest X-ray tube voltage at that current | 200 mA, 125 kV | - | - | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Minimum X-ray tube current and lowest X-ray tube voltage at that current | 10 mA, 40 kV | - | - | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | The lowest current time product (loading factors at the lowest current time product) | 0.007 mAs (2 mA, 3.5 ms) | - | - | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Note 1: Normal usage is 120 kV. For low load fluoroscopy, 125 kV is possible. |  | - | - | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | General note: Fluoroscopy is available in pulsed fluoroscopy mode only. |  | - | - | Jonne Moeskops
x | x |  | x |  |  | x | x |  |  |  | x | x
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Technical Data MRC 200 GS 0407 Tube | Tag: US.TechData.X-rayTube_MRC 200_GS_0407 | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Focal spot size and loadability | 0.4 mm nominal focal spot with maximal 30 kW loadability 0.7 mm nominal focal spot with maximal 65 kW loadability Based on 250 W anode reference power | D000416045 | TRS1103; TRS1104 | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Grid-switched pulsed fluoroscopy | yes | D000416045 | TRS1114 | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Fluoro power for 10 minutes | 4,500 W | 989000086501_100_000_00_1 | page 29 | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Fluoro power for 20 minutes | 3,500 W | 989000086501_100_000_00_1 | page 29 | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Maximum X-ray field with SID = 100 | 35 x 35 cm (13.8 x 13.8 inch) | D000416045 | TRS1614 | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Maximum X-ray field with SID = 120 | 42 x 42 cm (16.5 x 16.5 inch) | D000416045 | TRS1614 (calculated) | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Maximum X-ray field with SID = 70 | 24.5 x 24.5 cm (9.6 x 9.6 inch) | D000416045 | TRS1614 (calculated) | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Maximum anode cooling rate | 910 kHU/min | 989000086501_100_000_00_1 | page 96 | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Max. anode heat storage | 2.4 MHU | D000416045 | TRS1323 | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Max. assembly heat storage | 5.4 MHU | D000416045 | TRS1325 | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Anode heat dissipation | 11,000 W | 989000086501_100_000_00_1 | page 96 | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Continuous anode heat dissipation | 3,200 W | 989000086501_100_000_00_1 | page 96 | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Max. assembly continuous heat dissipation | 3,500 W | 989000086501_100_000_00_1 | page 29 | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | - | 3.2 X-ray tube | Extra pre-filtration | SpectraBeam filters: 0.1, 0.4, 0.9 mm Cu and 1 mm Al backing. | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | - | 3.2 X-ray tube | Cooling liquid | Oil cooled X-ray tube with thermal safety switch | 989000086501_100_000_00_1 | page 32 | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Anode target angle | 11° | D000416045 | TRS1121 | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | - | 3.2 X-ray tube | Anode cooling method | Direct anode oil cooling system with 200 mm anode diameter | 989000086501_100_000_00_1 | page 100 | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Tube Output Power |  | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Exposure |  | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Highest electric power (X-ray tube current and X-ray tube voltage resulting in highest electric power) | 65 kW (125 kV, 520 mA) | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Highest electric output power at 100kV, 0.1s (X-ray tube current) | 65 kW (650 mA) | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Maximum X-ray tube voltage and highest X-ray tube current at that voltage | 125 kV, 520 mA | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Maximum X-ray tube current and highest X-ray tube voltage at that current | 813 mA, 80 kV | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Minimum X-ray tube current and lowest X-ray tube voltage at that current | 10 mA, 40 kV | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | The lowest current time product (loading factors at the lowest current time product) | 0.1 mAs (10 mA, 10 ms) | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Fluoroscopy with Grid Switch |  | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Highest electric power (X-ray tube current and X-ray tube voltage resulting in highest electric power) | 20 kW (125 kV, 160 mA) | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Highest electric output power at 100kV, 0.1s (X-ray tube current) | 16 kW (160 mA) | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Maximum X-ray tube voltage and highest X-ray tube current at that voltage | 125 kV1, 160 mA | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Minimum X-ray tube voltage and lowest X-ray tube current at that voltage | 40 kV, 10 mA | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Maximum X-ray tube current and highest X-ray tube voltage at that current | 160 mA, 125 kV | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Minimum X-ray tube current and lowest X-ray tube voltage at that current | 10 mA, 40 kV | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | The lowest current time product (loading factors at the lowest current time product) | 0.007 mAs (2 mA, 3.5 ms) | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | Note 1: Normal usage is 120 kV. For low load fluoroscopy, 125 kV is possible. |  | - | - | Jonne Moeskops
X-ray generation | - | x | - | x | x | - | - | x | x | - | x | x | 1 | 3.2 X-ray tube | General note: Fluoroscopy is available in pulsed fluoroscopy mode only. |  | - | - | Jonne Moeskops
x | x |  | x |  |  | x | x |  |  |  | x | x
X-ray generation | x | - | x | - | - | x | x | - | - | x | x | x | 1 | 3.2 X-ray tube | Technical Data MRC 200 GS 0508 Tube | Tag: US.TechData.X-rayTube_MRC 200_GS_0508 | - | - | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Focal spot size and loadability | 0.5 mm nominal focal spot with maximal 45 kW loadability 0.8 mm nominal focal spot with maximal 85 kW loadability Based on 250 W anode reference power | D000416042 | TRS1103; TRS1104 | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Grid-switched pulsed fluoroscopy | yes | D000416042 | TRS1114 | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Fluoro power for 10 minutes | - | - | - | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Fluoro power for 20 minutes | - | - | - | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Maximum X-ray field with SID = 100 | 28 x 28 cm (11.0 x 11.0 inch) | D000416042 | TRS1614 | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Maximum X-ray field with SID = 120 | 33,6 x 33,6 cm (13.2 x 13.2 inch) | D000416042 | TRS1614 (calculated) | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Maximum X-ray field with SID = 70 | 19,6 x 19,6 cm (7.7 x 7.7 inch) | D000416042 | TRS1614 (calculated) | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Maximum anode cooling rate | 910 kHU/min | 989000003692_100_000_00_01 | page 100 | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Max. anode heat storage | 2.4 MHU | D000416042 | TRS1323 | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Max. assembly heat storage | 5.4 MHU | D000416042 | TRS1325 | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Anode heat dissipation | 11,000 W | 989000003692_100_000_00_01 | page 100 | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Continuous anode heat dissipation | 3,200 W | 989000003692_100_000_00_01 | page 100 | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Max. assembly continuous heat dissipation | 3,500 W | 989000003692_100_000_00_01 | page 54 | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | - | 3.2 X-ray tube | Extra pre-filtration | SpectraBeam filters: 0.1, 0.4, 0.9 mm Cu and 1 mm Al backing. | - | - | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | - | 3.2 X-ray tube | Cooling liquid | Oil cooled X-ray tube with thermal safety switch | 989000003692_100_000_00_01 | page 47, page 55 | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Anode target angle | 9° | D000416042 | TRS1121 | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | - | 3.2 X-ray tube | Anode cooling method | Direct anode oil cooling system with 200 mm anode diameter | 989000003692_100_000_00_01 | page 92 | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Tube Output Power |  | - | - | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Exposure |  | - | - | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Highest electric power (X-ray tube current and X-ray tube voltage resulting in highest electric power) | 85 kW (125 kV, 680 mA) | - | - | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Highest electric output power at 100kV, 0.1s (X-ray tube current) | 85 kW (850 mA) | - | - | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Maximum X-ray tube voltage and highest X-ray tube current at that voltage | 125 kV, 680 mA | - | - | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Maximum X-ray tube current and highest X-ray tube voltage at that current | 1063 mA, 80 kV (The maximum tube current cannot be reached with the current system configuration) | - | - | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Minimum X-ray tube current and lowest X-ray tube voltage at that current | 10 mA, 40 kV | - | - | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | The lowest current time product (loading factors at the lowest current time product) | 0.1 mAs (10 mA, 10 ms) | - | - | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Fluoroscopy with Grid Switch |  | - | - | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Highest electric power (X-ray tube current and X-ray tube voltage resulting in highest electric power) | 25 kW (125 kV, 200 mA) | - | - | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Highest electric output power at 100kV, 0.1s (X-ray tube current) | 20 kW (200 mA) | - | - | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Maximum X-ray tube voltage and highest X-ray tube current at that voltage | 125 kV1, 200 mA | - | - | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Minimum X-ray tube voltage and lowest X-ray tube current at that voltage | 40 kV, 10 mA | - | - | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Maximum X-ray tube current and highest X-ray tube voltage at that current | 200 mA, 125 kV | - | - | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Minimum X-ray tube current and lowest X-ray tube voltage at that current | 10 mA, 40 kV | - | - | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | The lowest current time product (loading factors at the lowest current time product) | 0.007 mAs (2 mA, 3.5 ms) | - | - | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | Note 1: Normal usage is 120 kV. For low load fluoroscopy, 125 kV is possible. |  | - | - | Jonne Moeskops
X-ray generation | x | - | x | - | - | x | x | - | - | - | x | x | 1 | 3.2 X-ray tube | General note: Fluoroscopy is available in pulsed fluoroscopy mode only. |  | - | - | Jonne Moeskops
x | x |  | x |  |  | x | x |  |  |  | x | x
X-ray generation | - | - | - | - | - | - | - | - | - | - | - | - | 1 | - | Accuracy of Dosimetric Indications and Automatic Control System |  | - | - | Jos van Tongeren
X-ray generation | - | - | - | - | - | - | - | - | - | - | - | - | 1 | - | Dosimetric Indication |  | - | - | Jos van Tongeren
X-ray generation | - | - | - | - | - | - | - | - | - | - | - | - | 1 | - | Accuracy of the reference air kerma | ±35% (above 100 mGy) | - | - | Jos van Tongeren
X-ray generation | - | - | - | - | - | - | - | - | - | - | - | - | 1 | - | Accuracy of reference air kerma rate | ±35% (above 6 mGy/min) | - | - | Jos van Tongeren
X-ray generation | - | - | - | - | - | - | - | - | - | - | - | - | 1 | - | Accuracy of cumulative dose area product | ±35% (above 2.5 Gy∙cm2) | - | - | Jos van Tongeren
X-ray generation | - | - | - | - | - | - | - | - | - | - | - | - | 1 | - | Coefficient of variation of the automatic control system | <0.05 | - | - | Jos van Tongeren
X-ray generation | - | - | - | - | - | - | - | - | - | - | - | - | 1 | - | Calculation method | The dosimetric indications for reference air kerma (rate) and cumulative dose area product are calculated using the acquisition parameters and a calibration procedure with a reference dose meter in the equipment. When the optional dose area product meter is used, the dose indications come directly from this instrument. | - | - | Jos van Tongeren
X-ray generation | - | - | - | - | - | - | - | - | - | - | - | - | 1 | - | Unit of Display | The unit of display for dose area product is configurable. | - | - | Jos van Tongeren
X-ray generation | - | - | - | - | - | - | - | - | - | - | - | - | 1 | - |  |  | - | - | Jos van Tongeren
x | x |  | x |  |  | x | x |  |  |  | x | x
X-ray generation | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 3.3 Collimator | Nicol V3 Collimator | Tag: US.TechData.NicolCollimator | DHF106048 | - | Ishaq Arman
X-ray generation | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 3.3 Collimator | Minimum field size | <2x2 mm at 1 m source-to-image distance | DHF106048 | - | Ishaq Arman
X-ray generation | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 3.3 Collimator | Inherent filtration (without spectral filter) | ≤0.1 mm Al equivalent at 75 kV | DHF106048 | - | Ishaq Arman
X-ray generation | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 3.3 Collimator | Maximum shutter operational speed at 1 m source-to-image distance | 20 cm/s | DHF106048 | - | Ishaq Arman
X-ray generation | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 3.3 Collimator | Maximum wedge operational speed at 1 m source-to-image distance | 10 cm/s | DHF106048 | - | Ishaq Arman
X-ray generation | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 3.3 Collimator | Maximum wedge operational rotational speed | 90 degrees/s | DHF106048 | - | Ishaq Arman
x | x |  | x |  |  | x | x |  |  |  | x | x
X-ray generation | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 3.3 Collimator | Nicol V4 Collimator | Tag: US.TechData.NicolCollimator | DHF202110: Nicol v4 | - | Ishaq Arman
X-ray generation | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 3.3 Collimator | Minimum field size | <2x2 mm at 1 m source-to-image distance | DHF202110: Nicol v4 | - | Ishaq Arman
X-ray generation | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 3.3 Collimator | Inherent filtration (without spectral filter) | ≤0.1 mm Al equivalent at 75 kV | DHF202110: Nicol v4 | - | Ishaq Arman
X-ray generation | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 3.3 Collimator | Maximum shutter operational speed at 1 m source-to-image distance | 20 cm/s | DHF202110: Nicol v4 | - | Ishaq Arman
X-ray generation | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 3.3 Collimator | Maximum wedge operational speed at 1 m source-to-image distance | 10 cm/s | DHF202110: Nicol v4 | - | Ishaq Arman
X-ray generation | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 3.3 Collimator | Maximum wedge operational rotational speed | 90 degrees/s | DHF202110: Nicol v4 | - | Ishaq Arman
X-ray generation | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 3.3 Collimator | Maximum collimator rotational speed (FlexArm option only) | 45 degrees/s | DHF207220: Nicol v4 | - | Ishaq Arman
x | x |  | x |  |  | x | x |  |  |  | x | x
X-ray generation | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 3.4 Anti-Scatter Grid | Anti-Scatter Grid FD12 | Tag: US.TechData.GridFD12 | - | - | -
X-ray generation | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 3.4 Anti-Scatter Grid | Lines/cm | 74 lines/cm | DHF217653 | - | -
X-ray generation | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 3.4 Anti-Scatter Grid | Ratio | 14 | DHF217653 | - | -
X-ray generation | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 3.4 Anti-Scatter Grid | focal Spot distance | 105 cm | DHF217653 | - | -
x | x |  | x |  |  | x | x |  |  |  | x | x
X-ray generation | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 3.4 Anti-Scatter Grid | Anti-Scatter Grid FD15 | Tag: US.TechData.GridFD15 | - | - | -
X-ray generation | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 3.4 Anti-Scatter Grid | Lines/cm | 70 lines/cm | DHF174368 | - | -
X-ray generation | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 3.4 Anti-Scatter Grid | Ratio | 13 | DHF174368 | - | -
X-ray generation | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 3.4 Anti-Scatter Grid | focal Spot distance | 100 cm | DHF174368 | - | -
x | x |  | x |  |  | x | x |  |  |  | x | x
X-ray generation | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 3.4 Anti-Scatter Grid | Anti-Scatter Grid FD20 | Tag: US.TechData.GridFD20 | - | - | -
X-ray generation | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 3.4 Anti-Scatter Grid | Lines/cm | 44 lines/cm | DHF144845 | - | -
X-ray generation | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 3.4 Anti-Scatter Grid | Ratio | 12 | DHF144845 | - | -
X-ray generation | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 3.4 Anti-Scatter Grid | focal Spot distance | 105 cm | DHF144845 | - | -
x | x | x | x | x | x | x | x | x | x | x | x | x
Imaging | - | x | - | - | - | - | - | - | - | - | - | x | 1 | 4.2 Dynamic Flat Detector | 15 Inch Detector | Tag: US.TechData.15InchDetectore.PX2630Sv | CE103412 | - | Jonne Moeskops
Imaging | - | x | - | - | - | - | - | - | - | - | - | x | 1 | 4.2 Dynamic Flat Detector | Size of detector housing | 58 cm (22.8 inch) diagonal | - | - | Maikel van Eekelen
Imaging | - | x | - | - | - | - | - | - | - | - | - | x | 1 | 4.2 Dynamic Flat Detector | Detector dimensions | 33.8 x 39.2 cm | DHF141218 | - | Maikel van Eekelen
Imaging | - | x | - | - | - | - | - | - | - | - | - | x | 1 | 4.2 Dynamic Flat Detector | Maximum field of view | 39 cm (15.2 inch) diagonal | D000764665 | - | Erik Oerlemans
Imaging | - | x | - | - | - | - | - | - | - | - | - | x | 1 | 4.2 Dynamic Flat Detector | X-ray sensitive area | 1420 x 1560 pixels, 261 mm x 287 mm | DHF141218 | - | -
Imaging | - | x | - | - | - | - | - | - | - | - | - | x | 1 | 4.2 Dynamic Flat Detector | Detector zoom fields | 39, 37, 31, 27, 22, 19, 15 cm (15, 14, 13, 10, 8, 7, 6 inch) diagonal square formats | D000764665 | - | Erik Oerlemans
Imaging | - | x | - | - | - | - | - | - | - | - | - | x | 1 | 4.2 Dynamic Flat Detector | Pixel pitch | 184 μm x 184 μm | DHF137418 | [PTS-REQ-0600] | Jonne Moeskops
Imaging | - | x | - | - | - | - | - | - | - | - | - | x | 1 | 4.2 Dynamic Flat Detector | Detector bit depth | 16 bits | DHF137418 | [PTS-REQ-0690] | Jonne Moeskops
Imaging | - | x | - | - | - | - | - | - | - | - | - | x | 1 | 4.2 Dynamic Flat Detector | Nyquist frequency | 2.7 lp/mm | DHF137418 | =1/(2*pixel pitch) = 1/ (368 μm) = 0,002717 lp/μm => 2,7 lp/mm | Jonne Moeskops
Imaging | - | x | - | - | - | - | - | - | - | - | - | x | 1 | 4.2 Dynamic Flat Detector | DQE (0) | 70% at 0 lp/mm | DHF137418 | [PTS-REQ-1450] | Jonne Moeskops
Imaging | - | x | - | - | - | - | - | - | - | - | - | x | 1 | 4.2 Dynamic Flat Detector | MTF at 1 lp/mm | 59% (typical) | DHF137418 | [PTS-REQ-1490] | Jonne Moeskops
Imaging | - | - | - | x | x | - | - | x | x | - | x | x | 1 | 4.2 Dynamic Flat Detector | 20 Inch Detector | Tag: US.TechData.20InchDetector.PX3040F | CE104368 | - | Jonne Moeskops
Imaging | - | - | - | x | x | - | - | x | x | - | x | x | 1 | 4.2 Dynamic Flat Detector | Size of detector housing | 67 cm (26 inch) diagonal, including BodyGuard | - | - | Maikel van Eekelen
Imaging | - | - | - | x | x | - | - | x | x | - | x | x | 1 | 4.2 Dynamic Flat Detector | Detector dimensions | 47.2 x 36.0 cm | DHF160502 | - | Maikel van Eekelen
Imaging | - | - | - | x | x | - | - | x | x | - | x | x | 1 | 4.2 Dynamic Flat Detector | Maximum field of view | 48 cm (19 inch) diagonal | D000764665 | - | Erik Oerlemans
Imaging | - | - | - | x | x | - | - | x | x | - | x | x | 1 | 4.2 Dynamic Flat Detector | X-ray sensitive area | 1904 x 2586 pixels, 293.2 mm x 398.2 mm | DHF160502 | - | Erik Oerlemans
Imaging | - | - | - | x | x | - | - | x | x | - | x | x | 1 | 4.2 Dynamic Flat Detector | Detector zoom fields | 48, 42, 37, 31, 27, 22, 19, 15 cm (19, 17, 14.4, 13, 10.5, 8, 7, 6 inch) diagonal formats | D000764665 | - | Erik Oerlemans
Imaging | - | - | - | x | x | - | - | x | x | - | x | x | 1 | 4.2 Dynamic Flat Detector | Pixel pitch | 154 μm x 154 μm | DHF160503 | [PTS-REQ-0600] | Jonne Moeskops
Imaging | - | - | - | x | x | - | - | x | x | - | x | x | 1 | 4.2 Dynamic Flat Detector | Detector bit depth | 16 bits | DHF160503 | [PTS-REQ-0690] | Jonne Moeskops
Imaging | - | - | - | x | x | - | - | x | x | - | x | x | 1 | 4.2 Dynamic Flat Detector | Nyquist frequency | 3.2 lp/mm | - | =1/(2*pixel pitch) = 1/ (308 μm) = 0,003246 lp/μm => 3,2 lp/mm | Jonne Moeskops
Imaging | - | - | - | x | x | - | - | x | x | - | x | x | 1 | 4.2 Dynamic Flat Detector | DQE (0) | 77% at 0 lp/mm | DHF160503 | [PTS-REQ-1450] | Jonne Moeskops
Imaging | - | - | - | x | x | - | - | x | x | - | x | x | 1 | 4.2 Dynamic Flat Detector | MTF at 1 lp/mm | 59% (typical) | DHF160503 | [PTS-REQ-1490] | Jonne Moeskops
Imaging | x | - | x | - | - | x | x | - | - | x | x | - | 1 | 4.2 Dynamic Flat Detector | 12 Inch Detector | Tag: US.TechData.12InchDetector.PX2121S/CV | CE104975 | - | Jonne Moeskops
Imaging | x | - | x | - | - | x | x | - | - | x | x | - | 1 | 4.2 Dynamic Flat Detector | Size of detector housing | 47 cm (18 inch) diagonal including BodyGuard | - | - | Maikel van Eekelen
Imaging | x | - | x | - | - | x | x | - | - | x | x | - | 1 | 4.2 Dynamic Flat Detector | Detector dimensions | 28.8 x 28.3 cm | DHF209021 | - | Maikel van Eekelen
Imaging | x | - | x | - | - | x | x | - | - | x | x | - | 1 | 4.2 Dynamic Flat Detector | Maximum field of view | 30 cm (12 inch) diagonal | D000764665 | - | Erik Oerlemans
Imaging | x | - | x | - | - | x | x | - | - | x | x | - | 1 | 4.2 Dynamic Flat Detector | X-ray sensitive area | 1344 x 1344 pixels (207 mm x 207 mm) | DHF209021 | - | Erik Oerlemans
Imaging | x | - | x | - | - | x | x | - | - | x | x | - | 1 | 4.2 Dynamic Flat Detector | Detector zoom fields | 30, 27, 22, 19, 15 cm (12, 11, 8, 7, 6 inch) diagonal square formats | D000764665 | - | Erik Oerlemans
Imaging | x | - | x | - | - | x | x | - | - | x | x | - | 1 | 4.2 Dynamic Flat Detector | Pixel pitch | 154 μm x 154 μm | DHF168424 | [PTS-REQ-0600] | Jonne Moeskops
Imaging | x | - | x | - | - | x | x | - | - | x | x | - | 1 | 4.2 Dynamic Flat Detector | Detector bit depth | 16 bits | DHF168424 | [PTS-REQ-0690] | Jonne Moeskops
Imaging | x | - | x | - | - | x | x | - | - | x | x | - | 1 | 4.2 Dynamic Flat Detector | Nyquist frequency | 3.2 lp/mm | - | =1/(2*pixel pitch) = 1/ (308 μm) = 0,003246 lp/μm => 3,2 lp/mm | Jonne Moeskops
Imaging | x | - | x | - | - | x | x | - | - | x | x | - | 1 | 4.2 Dynamic Flat Detector | DQE (0) | 77% at 0 lp/mm | DHF168424 | [PTS-REQ-1450] | Jonne Moeskops
Imaging | x | - | x | - | - | x | x | - | - | x | x | - | 1 | 4.2 Dynamic Flat Detector | MTF at 1 lp/mm | 59% (typical) | DHF168424 | [PTS-REQ-1490] | Jonne Moeskops
Imaging | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 4.3 Fluoroscopy | Fluoroscopy | Tag: US.TechData.Fluoroscopy | - | - | -
Imaging | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 4.3 Fluoroscopy | Extra pre-filtration | SpectraBeam filters: 0.1, 0.4, 0.9 mm Cu and 1 mm Al backing. | - | - | Jan Jans
Imaging | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 4.3 Fluoroscopy | Fluoroscopy image processing | Smart Recursive filtering, localized contrastadaptive contour enhancement and Clarity-IQ algorithm | - | - | Ben Gevers
Imaging | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 4.3 Fluoroscopy | Pulse rates | 30, 25, 20, 15, 12.5, 10, 7.5, 6.25, 5, 3.75, 3.125, 2.5, 1.875, 1.25, 1.0, 0.625, 0.5 images/sec. | SRS | SRS.Allura.Func.FluoFlavours | Jeroen Helwegen
Imaging | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 4.3 Fluoroscopy | Frame grabbing of static fluoroscopy images | Yes | SRS | SRS.Allura.Func.StoreFluoro | Jeroen Helwegen
Imaging | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 4.3 Fluoroscopy | Fluoroscopy storage | Default storage of the up to 2000 images, based on PulsedFluoFrameSpeed (EPX defined) x FluoStoreLength (EPX defined)) of fluoroscopy for reference or archiving | UID | UID.Allura.Acquisition.StoreFluoro | Jeroen Helwegen
Imaging | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 4.3 Fluoroscopy | Grid-switched pulsed fluoroscopy | Yes | - | - | Jeroen Helwegen
Imaging | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 4.4 Digital Acquisition | Digital Acquisition | Tag: US.TechData.DigitalAcquisition | - | - | -
Imaging | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 4.4 Digital Acquisition | Standard configuration | 0,5 to 12 images/sec | SRS, UID | SRS.Allura.Func.ExposureFramespeed + UID.Allura.Acquisition.ImageGen.CineExposures | Jeroen Helwegen
Imaging | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 4.4 Digital Acquisition | Image storage | 50.000 images (based on 1k2) | SRS | SRS.Allura.Conf.ImageStorageOptions.xs-2 | Jeroen Helwegen
Imaging | x | x | x | x | x | x | x | x | x | x | x | x | 1 | 4.4 Digital Acquisition | Storage extension (optional) | 100.000 images (based on 1k2) | SRS | SRS.Allura.Conf.ImageStorageOptions.xs-3 | Jeroen Helwegen
x | x | x | x | x | x | x | x | x | x | x | x | x
Viewing | x | x | x | x | x | x | x | x | x | x | x | x | - | 5.1 Monitors | 27" LCD ER | Tag: US.TechData.Mon27LCD | DHF225824 | - | Marc Arends
Viewing | x | x | x | x | x | x | x | x | x | x | x | x | - | 5.1 Monitors | Size | 27 inch high brightness color TFT-LCD display | DHF225824 | Par 2.3 Composition | Marc Arends
Viewing | x | x | x | x | x | x | x | x | x | x | x | x | - | 5.1 Monitors | Format | Native format 1920x1080 Full HD | DHF225824 | Par 2.3 Composition | Marc Arends
Viewing | x | x | x | x | x | x | x | x | x | x | x | x | - | 5.1 Monitors | Wide viewing angle | Wide viewing angle (approx. 178 degrees) | DHF225824 | US.WidescreenMonitors.Performance.FoS | Marc Arends
Viewing | x | x | x | x | x | x | x | x | x | x | x | x | - | 5.1 Monitors | High Brightness | High brightness (max 650 Cd/m2, default 400 Cd/m2) | DHF225824 DHF330332 | US.WidescreenMonitors.Performance.FoS DHF330332 PPA, 16-DMR, Datasheets, K5902101_01_UserGuide.pdf | Marc Arends
Viewing | - | - | - | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | Contrast ratio | 1000:1 (typical) | DHF225824 | US.WidescreenMonitors.Performance.FoS | Marc Arends
Viewing | x | x | x | x | x | x | x | x | x | x | x | x | - | 5.1 Monitors | Luminance stability | Long term luminance stability through backlight stabilization circuit | DHF225824 | US.WidescreenMonitors.Performance.FoS | Marc Arends
Viewing | x | x | x | x | x | x | x | x | x | x | x | x | - | 5.1 Monitors | Brightness control | Automatic brightness control with backlight sensor | DHF225824 DHF330332 | US.WidescreenMonitors.Performance.FoS DHF330332 PPA,03_Design records-specifications,161-991932282074 MMEC-2127_PS27FHDER_19042019.pdf | Marc Arends
Viewing | x | x | x | x | x | x | x | x | x | x | x | x | - | 5.1 Monitors | Control functions | Control functions on side | DHF225824 | US.WidescreenMonitors.DesignConstraints.OSD | Marc Arends
Viewing | x | x | x | x | x | x | x | x | x | x | x | x | - | 5.1 Monitors | On-screen display | Yes | DHF225824 | US.WidescreenMonitors.DesignConstraints.OSD | Marc Arends
Viewing | x | x | x | x | x | x | x | x | x | x | x | x | - | 5.1 Monitors | Lookup table | Internal selectable lookup table for gray-scale transfer function, including DICOM | DHF225824 | US.WidescreenMonitors.Performance.GreyscaleTransferFunction | Marc Arends
Viewing | x | x | x | x | x | x | x | x | x | x | x | x | - | 5.1 Monitors | Power supply | Internal power supply (100-240 VAC) | DHF225824 | US.WidescreenMonitors.Performance.Resource.Power | Marc Arends
Viewing | x | x | x | x | x | x | x | x | x | x | x | x | - | 5.1 Monitors | Protection screen | Full protective screen Ingress Protection: IP21 | DHF225824 | US.WidescreenMonitors.Functional.ProtectiveScreen | Marc Arends
Viewing | - | - | - | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | Connectivity | N.A. | DHF225824 | US.WidescreenMonitors.DesignConstraints.USB | Marc Arends
Viewing | x | x | x | x | x | x | x | x | x | x | x | x | - | 5.1 Monitors | 24" LCD CR | Tag: US.TechData.Mon24LCD | DHF225824 | - | Marc Arends
Viewing | - | - | - | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | Size | 24 inch high brightness color TFT-LCD display | DHF225824 | Par 2.3 Composition | Marc Arends
Viewing | - | - | - | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | Format | Native format 1920x1080 Full HD | DHF225824 | Par 2.3 Composition | Marc Arends
Viewing | - | - | - | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | Wide viewing angle | Wide viewing angle (approx. 178 degrees) | DHF225824 | US.WidescreenMonitors.Performance.FoS | Marc Arends
Viewing | - | - | - | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | High Brightness | High brightness (max 400 Cd/m2, default 350 Cd/m2) | DHF225824 D000845262 | US.WidescreenMonitors.Performance.FoS Philips RUP2433AMI User Guide_Rev.A | Marc Arends
Viewing | - | - | - | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | Contrast ratio | 1000:1 (typical) | DHF225824 DHF352137 | US.WidescreenMonitors.Performance.FoS PSW 24" Harmonized monitor, Evidences, 01.Design Input for APQP, GRN2433AMI_1A1A24210112200_Engineering Specification Rev. K | Marc Arends
Viewing | - | - | - | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | Luminance stability | Long term luminance stability through backlight stabilization circuit | DHF225824 DHF352137 | US.WidescreenMonitors.Performance.FoS PSW 24" Harmonized monitor, Evidences, 01.Design Input for APQP, GRN2433AMI_1A1A24210112200_Engineering Specification Rev. K | Marc Arends
Viewing | - | - | - | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | Brightness control | Automatic brightness control with backlight sensor | DHF225824 DHF352137 | US.WidescreenMonitors.Performance.FoS PSW 24" Harmonized monitor, Evidences, 01.Design Input for APQP, GRN2433AMI_1A1A24210112200_Engineering Specification Rev. K | Marc Arends
Viewing | - | - | - | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | Control functions | Control functions on side | DHF225824 | US.WidescreenMonitors.DesignConstraints.OSD | Marc Arends
Viewing | - | - | - | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | On-screen display | Yes | DHF225824 | US.WidescreenMonitors.DesignConstraints.OSD | Marc Arends
Viewing | - | - | - | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | Lookup table | Internal selectable lookup table for gray-scale transfer function, including DICOM | DHF225824 | US.WidescreenMonitors.Performance.GreyscaleTransferFunction | Marc Arends
Viewing | - | - | - | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | Power supply | Internal power supply (100-240 VAC) | DHF225824 | US.WidescreenMonitors.Performance.Resource.Power | Marc Arends
Viewing | - | - | - | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | Protection screen | N.A. | DHF225824 | US.WidescreenMonitors.Functional.ProtectiveScreen | Marc Arends
Viewing | - | - | - | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | Connectivity | Integrated USB hub with 2 USB ports on side | DHF225824 | US.WidescreenMonitors.DesignConstraints.USB | Marc Arends
Viewing | - | - | x | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | 32" LCD ER (optional requires 3rd party boom) | Tag: US.TechData.Mon32LCD | DHF225824 | - | Marc Arends
Viewing | - | - | x | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | Size | 32 inch high brightness color TFT-LCD display | DHF225824 | Par 2.3 Composition | Marc Arends
Viewing | - | - | x | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | Format | Native format 1920x1080 Full HD | DHF225824 | Par 2.3 Composition | Marc Arends
Viewing | - | - | x | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | Wide viewing angle | Wide viewing angle (approx. 178 degrees) | DHF225824 | US.WidescreenMonitors.Performance.FoS | Marc Arends
Viewing | - | - | x | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | High Brightness | High brightness (max 500 Cd/m2, default 400 Cd/m2) | DHF225824 DHF310526 | US.WidescreenMonitors.Performance.FoS; DHF310526 PPA, 16-DMR, K5902107_01AB_UserGuide.pdf | Marc Arends
Viewing | - | - | - | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | Contrast ratio | 1300:1 (typical) | DHF310526 | DHF310526 PPA, 16-DMR, K5902107_01AB_UserGuide.pdf | Marc Arends
Viewing | - | - | x | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | Luminance stability | Long term luminance stability through backlight stabilization circuit | DHF225824 | US.WidescreenMonitors.Performance.FoS | Marc Arends
Viewing | - | - | x | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | Brightness control | Automatic brightness control with backlight sensor | DHF225824 DHF310526 | US.WidescreenMonitors.Performance.FoS DHF310526 PPA, 03_Design records, 161-991932282093 MMEC-2132_PS32FHDER_16032018.pdf | Marc Arends
Viewing | - | - | x | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | Control functions | Control functions on side | DHF225824 | US.WidescreenMonitors.DesignConstraints.OSD | Marc Arends
Viewing | - | - | x | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | On-screen display | Yes | DHF225824 | US.WidescreenMonitors.DesignConstraints.OSD | Marc Arends
Viewing | - | - | x | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | Lookup table | Internal selectable lookup table for gray-scale transfer function, including DICOM | DHF225824 | US.WidescreenMonitors.Performance.GreyscaleTransferFunction | Marc Arends
Viewing | - | - | x | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | Power supply | Internal power supply (100-240 VAC) | DHF225824 | US.WidescreenMonitors.Performance.Resource.Power | Marc Arends
Viewing | - | - | x | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | Protection screen | Full protective screen Ingress Protection: IP21 | DHF225824 | US.FlexVisionMonitor.Functional.ProtectiveScreen US.FlexVisionMonitor.DesignConstraints.IngressProtection | Marc Arends
Viewing | - | - | - | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | Connectivity | N.A. | DHF225824 | US.WidescreenMonitors.DesignConstraints.USB | Marc Arends
Viewing | x | x | x | x | x | x | x | x | x | x | x | x | - | 5.1 Monitors | Optional monitors in Exam Room: | - | - | - | Marc Arends
Viewing | x | x | x | x | x | x | x | x | x | x | x | x | - | 5.1 Monitors | FlexVision XL | Tag: US.TechData.FlexVisionXL | - | - | Marc Arends
Viewing | x | x | x | x | x | x | x | x | x | x | x | x | - | 5.1 Monitors | Size | 58-inch, 8 Megapixel color LCD | DHF233619 | Par 2.2 Responsibilities | Marc Arends
Viewing | x | x | x | x | x | x | x | x | x | x | x | x | - | 5.1 Monitors | Format | Native resolution: 3840x2160 | DHF233619 | US.FlexVisionMonitor.Functional.Video | Marc Arends
Viewing | - | - | - | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | Wide viewing angle | Approx. 176 degrees | DHF233619 | US.FlexVisionMonitor.Performance.FoS | Marc Arends
Viewing | x | x | x | x | x | x | x | x | x | x | x | x | - | 5.1 Monitors | High Brightness | Max: 700 Cd/m2 (typical) stabilized: 400 Cd/m2 | DHF233619 DHF306564 | US.FlexVisionMonitor.Performance.FoS DHF384565, R5910201 User Manual CML 5883 | Marc Arends
Viewing | x | x | x | x | x | x | x | x | x | x | x | x | - | 5.1 Monitors | Contrast ratio | 4000:1 (typical) | DHF233619 | US.FlexVisionMonitor.Performance.FoS | Marc Arends
Viewing | - | - | - | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | Luminance stability | Long term luminance stability through backlight stabilization circuit | DHF233619 | US.FlexVisionMonitor.Performance.FoS | Marc Arends
Viewing | - | - | - | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | Brightness control | Automatic brightness control with backlight sensor | DHF233619 DHF306564 | US.FlexVisionMonitor.Performance.FoS PPA Evidence, 26-Norm Compliancy, E346057-D1018-CB-1-Original | Marc Arends
Viewing | - | - | - | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | Control functions | Control functions on bottom side | DHF233619 | US.FlexVisionMonitor.DesignConstraints.OSD | Marc Arends
Viewing | - | - | - | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | On-screen display | Yes | DHF233619 | US.FlexVisionMonitor.DesignConstraints.OSD | Marc Arends
Viewing | x | x | x | x | x | x | x | x | x | x | x | x | - | 5.1 Monitors | Lookup table | Internal selectable lookup table for gray-scale transfer function, including DICOM | DHF233619 | US.FlexVisionMonitor.Performance.GreyscaleTransferFunction | Marc Arends
Viewing | - | - | - | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | Power supply | Internal power supply (100-240 VAC) | DHF233619 DHF306564 | US.FlexVisionMonitor.Performance.Resource.Power DHF384565, R5910201 User Manual CML 5883 | Marc Arends
Viewing | x | x | x | x | x | x | x | x | x | x | x | x | - | 5.1 Monitors | Protection screen | Full protective screen Ingress Protection: IP21 | DHF233619 | US.FlexVisionMonitor.Functional.ProtectiveScreen US.FlexVisionMonitor.DesignConstraints.IngressProtection | Marc Arends
Viewing | - | - | - | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | Connectivity | N.A. | - | - | Marc Arends
Viewing | x | x | - | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | 19" LCD ER | Tag: US.TechData.Mon19LCD | - | - | Marc Arends
Viewing | x | x | - | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | Size | 19 inch high brightness color TFT-LCD display | DHF232737 | C15S Specification for Philips.pdf, par. 2.4 panel features | Marc Arends
Viewing | x | x | - | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | Format | Native format 1280x1024 SXGA | DHF232737 | C15S Specification for Philips.pdf, par. 2.4 panel features | Marc Arends
Viewing | x | x | - | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | Wide viewing angle | Wide viewing angle (approx. 178 degrees) | DHF232737 | C15S Specification for Philips.pdf, par. 2.5.3 View angle | Marc Arends
Viewing | x | x | - | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | High Brightness | High brightness (max 600 Cd/m2, default 400 Cd/m2) | D000781438, DHF232737 | PSW BigTide 300006525241, Evidence, 13 First Article Inspection, Attachment 1 - G190EAN01.5_ver.1_2018.05.25_Final Spec; PPA C15S User Manual Rev A1.pdf, Brightness | Marc Arends
Viewing | - | - | - | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | Contrast ratio | 1000:1 (typical) | D000781438 | PSW BigTide 300006525241, Evidence, 13 First Article Inspection, Attachment 1 - G190EAN01.5_ver.1_2018.05.25_Final Spec; PPA C15S User Manual Rev A1.pdf, Brightness | Marc Arends
Viewing | x | x | - | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | Luminance stability | Long term luminance stability through backlight stabilization circuit | DHF232737 | PPA C15S User Manual Rev A1.pdf; 1. Application, Fast backlight stability | Marc Arends
Viewing | x | x | - | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | Brightness control | Automatic brightness control with backlight sensor | DHF232737 | PPA C15S User Manual Rev A1.pdf; 1. Application, Fast backlight stability | Marc Arends
Viewing | x | x | - | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | Control functions | Control functions on front | DHF232737 | PPA C15S User Manual Rev A1.pdf; 6.3 OSD | Marc Arends
Viewing | x | x | - | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | On Screen Display | Yes | DHF232737 | PPA C15S User Manual Rev A1.pdf; 6.3 OSD menu | Marc Arends
Viewing | x | x | - | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | Lookup table | Internal selectable lookup table for gray-scale transfer function, including DICOM | DHF232737 | PPA C15S User Manual Rev A1.pdf; 1. Application, Embedded LUT | Marc Arends
Viewing | x | x | - | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | Power supply | Internal power supply (100-240 VAC) | DHF232737 | PPA C15S User Manual Rev A1.pdf; 8.2 Power supply | Marc Arends
Viewing | - | - | - | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | Protection screen | N.A. Ingress Protection IP20 | DHF232737 | PPA C15S User Manual Rev A1.pdf; 8.6 Mechanical Design | Marc Arends
Viewing | - | - | - | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | Connectivity | N.A. | - | - | Marc Arends
Viewing | - | - | - | - | - | - | - | - | - | x | x | x | - | 5.1 Monitors | Optional monitors in Control room: | - | - | - | Marc Arends
Viewing | - | - | - | - | - | - | - | - | - | x | x | x | - | 5.1 Monitors | Flexspot | Tag: US.TechData.FlexSpot | DHF341886 | - | Marc Arends
Viewing | - | - | - | - | - | - | - | - | - | x | x | x | - | 5.1 Monitors | Size | 27 inch color TFT-LCD display | DHF225824 | Par 2.3 Composition | Marc Arends
Viewing | - | - | - | - | - | - | - | - | - | x | x | x | - | 5.1 Monitors | Format | Native format 2560x1440 Quad HD | DHF225824 | Par 2.3 Composition | Marc Arends
Viewing | - | - | - | - | - | - | - | - | - | x | x | x | - | 5.1 Monitors | Wide viewing angle | Wide viewing angle (approx. 178 degrees) | DHF225824 | US.WidescreenMonitors.Performance.FoS | Marc Arends
Viewing | - | - | - | - | - | - | - | - | - | x | x | x | - | 5.1 Monitors | High Brightness | High brightness (max 500 Cd/m2, default 350 Cd/m2) | DHF225824 DHF352122 | US.WidescreenMonitors.Performance.FoS US.WidescreenMonitors.Performance.GreyscaleTransferFunction DHF352122, 27-inch monitor PSW Evidences, 01.Design input for APQP, GUP2732CMI_1A1C27180112200_Engineering Specification Rev. H.pdf | Marc Arends
Viewing | - | - | - | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | Contrast ratio | 1000:1 (typical) | DHF225824 DHF352122 | US.WidescreenMonitors.Performance.FoS DHF352122, 27-inch monitor PSW Evidences, 01.Design input for APQP, GUP2732CMI_1A1C27180112200_Engineering Specification Rev. H.pdf | Marc Arends
Viewing | - | - | - | - | - | - | - | - | - | x | x | x | - | 5.1 Monitors | Luminance stability | Long term luminance stability through backlight stabilization circuit | DHF225824 | US.WidescreenMonitors.Performance.FoS | Marc Arends
Viewing | - | - | - | - | - | - | - | - | - | x | x | x | - | 5.1 Monitors | Brightness control | Automatic brightness control with backlight sensor | DHF225824 DHF352122 | US.WidescreenMonitors.Performance.FoS DHF352122, 27-inch monitor PSW Evidences, 01.Design input for APQP, GUP2732CMI_1A1C27180112200_Engineering Specification Rev. H.pdf | Marc Arends
Viewing | - | - | - | - | - | - | - | - | - | x | x | x | - | 5.1 Monitors | Control functions | Control functions on side | DHF225824 | US.WidescreenMonitors.DesignConstraints.OSD | Marc Arends
Viewing | - | - | - | - | - | - | - | - | - | x | x | x | - | 5.1 Monitors | On-screen display | yes | DHF225824 | US.WidescreenMonitors.DesignConstraints.OSD | Marc Arends
Viewing | - | - | - | - | - | - | - | - | - | x | x | x | - | 5.1 Monitors | Lookup table | Internal selectable lookup table for gray-scale transfer function, including DICOM | DHF225824 | US.WidescreenMonitors.Performance.GreyscaleTransferFunction | Marc Arends
Viewing | - | - | - | - | - | - | - | - | - | x | x | x | - | 5.1 Monitors | Power supply | Internal power supply (100-240 VAC) | DHF225824 | US.WidescreenMonitors.Performance.Resource.Power | Marc Arends
Viewing | - | - | - | - | - | - | - | - | - | - | - | - | - | 5.1 Monitors | Protection screen | N.A. Ingress Protection IPX0 | DHF225824 | US.WidescreenMonitors.DesignConstraints.IngressProtection | Marc Arends
Viewing | - | - | - | - | - | - | - | - | - | x | x | x | - | 5.1 Monitors | Connectivity | Integrated USB hub with 2 USB ports on side | DHF225824 | US.WidescreenMonitors.DesignConstraints.USB | Marc Arends
Viewing | - | - | - | - | - | - | - | - | - | x | x | x | - | 5.1 Monitors | Additional Flexspot for DVI-Based Systems | This feature adds a second FlexSpot workspot with its own high resolution Quad HD (2560x1440) display as well as its own keyboard and mouse. • Up to 1 video source can be displayed at a time on the add-on FlexSpot display. • The X-ray status area with all X-ray details can be shown/hidden. | DHF341886 | US.WVD.DVI.Configurations.Workspots.ControlRoom US.WVD.DVI.FlexCompositor US.WVD.DVI.WorkspotControl | Marc Arends
Viewing | - | - | - | - | - | - | - | - | - | x | x | x | - | 5.1 Monitors | Additional Flexspot for IP-Based Systems | This feature adds a second FlexSpot workspot consisting of one or two Quad HD (2560x1440) high-resolution monitors to the primary FlexSpot workspot. It enables the user to show up to 8 video sources on the second FlexSpot workspot by combining up to 2 high resolution displays. Keyboard and mouse control is seamless across the 2 displays. | DHF299949 | US.WVD.IP.Configurations.Workspots.ControlRoom US.WVD.IP.WorkspotControl US.WVD.IP.FlexCompositor | Marc Arends
Viewing | - | - | - | - | - | - | - | - | - | x | x | x | - | 5.1 Monitors | 2nd display for FlexSpot | This feature adds a second Quad HD full (2560x1440) high-resolution monitor to the primary FlexSpot workspot. It enables the user to show up to 8 video sources on a single FlexSpot workspot by combining 2 high resolution displays. Keyboard and mouse control is seamless across the 2 displays. | DHF341886 | US.WVD.DVI.Configurations.Workspots.ControlRoom US.WVD.DVI.FlexCompositor | Marc Arends
x | x | x | x | x | x | x | x | x | x | x | x | x
System Infrastructure | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 6.1 Integration | Video Connection for DVI-Based Systems | Tag: US.TechData.VideoConnectionDVI |  |  | Marc Arends
System Infrastructure | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 6.1 Integration | Standard | DVI 1.0 | DHF220622 | IDS.DVI.VideoInputConnectivity.Digital.To.Dedicated_X-Ray_System_Monitor | Marc Arends
System Infrastructure | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 6.1 Integration | Connector type | DVI-I | DHF220622 | IDS.DVI.VideoInputConnectivity.Digital.To.Dedicated_X-Ray_System_Monitor | Marc Arends
System Infrastructure | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 6.1 Integration | Cable length (external equipment side) | 3 m DVI-I to DVI-I cable, 3 m VGA to DVI-I cable | DHF220622 | IDS.DVI.VideoInputConnectivity.Digital.To.Dedicated_X-Ray_System_Monitor | Marc Arends
System Infrastructure | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 6.1 Integration | Supported resolutions | Up to 1920 x 1200 x 60 Hz (WUXGA) | DHF220622 | IDS.DVI.VideoInputConnectivity.Digital.To.Dedicated_X-Ray_System_Monitor | Marc Arends
System Infrastructure | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 6.1 Integration | Supported clock frequencies | 25-165 MHz | DHF220622 | IDS Appendix B | Marc Arends
System Infrastructure | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 6.1 Integration | DVI lanes supported | 1 | DHF220622 | IDS.DVI.VideoInputConnectivity.Digital.To.Dedicated_X-Ray_System_Monitor | Marc Arends
System Infrastructure | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 6.1 Integration | Supported features | EDID / DDC2, Hot Plug Detect optional | DHF220622 | IDS Appendix B | Marc Arends
System Infrastructure | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 6.1 Integration | Video Connection for IP-Based Systems | Tag: US.TechData.VideoConnectionIP |  |  | Marc Arends
System Infrastructure | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 6.1 Integration | Standard | DVI 1.0 / DisplayPort 1.2 | DHF310478 | US.AVIDIS.Functional.AUXEA-DP-DVI | Marc Arends
System Infrastructure | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 6.1 Integration | Connector type | DVI-I / DisplayPort (only one connection is allowed at a time) | DHF310478 | US.AVIDIS.Functional.AUXEA-DP-DVI | Marc Arends
System Infrastructure | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 6.1 Integration | Cable length (external equipment side) | 3 m DVI-I to DVI-I cable, 3 m DP to DP cable | DHF220622 | IDS.VOIP.Video.Input2K | Marc Arends
System Infrastructure | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 6.1 Integration | Supported resolutions | Up to 1920 x 1200 x 60 Hz (WUXGA) | DHF310478 | US.AVIDIS.Functional.RestrictInputResolution | Marc Arends
System Infrastructure | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 6.1 Integration | Supported clock frequencies | 25-165 MHz | DHF220622 | IDS Appendix C | Marc Arends
System Infrastructure | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 6.1 Integration | DVI lanes supported | 1 | DHF310478 | US.AVIDIS.Functional.AUXEA-DP-DVI | Marc Arends
System Infrastructure | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 6.1 Integration | DP lanes supported | 4 lanes, SST | DHF310478 | US.AVIDIS.Functional.AUXEA-DP-DVI | Marc Arends
System Infrastructure | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 6.1 Integration | Supported features | EDID / DDC2, Hot Plug Detect optional | DHF220622 | IDS Appendix C | Marc Arends
System Infrastructure | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 6.1 Integration | Video Connection for 4K-Based Systems | Tag: US.TechData.VideoConnectionIP |  |  | Marc Arends
System Infrastructure | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 6.1 Integration | Standard | DisplayPort 1.2 | DHF310478 | US.AVIDIS.Functional.AUXEA-DP-4K | Marc Arends
System Infrastructure | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 6.1 Integration | Connector type | DisplayPort | DHF310478 | US.AVIDIS.Functional.AUXEA-DP-4K | Marc Arends
System Infrastructure | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 6.1 Integration | Cable length (external equipment side) | 3 m DP to DP cable | DHF220622 | IDS.VOIP.Video.Input4K | Marc Arends
System Infrastructure | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 6.1 Integration | Supported resolutions | Up to 3840 x 2160 x 60 Hz (UHD) | DHF220622 | IDS.VOIP.Video.Input4K | Marc Arends
System Infrastructure | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 6.1 Integration | Supported clock frequencies | 25-600 MHz | DHF220622 | IDS Appendix C | Marc Arends
System Infrastructure | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 6.1 Integration | DP lanes supported | 4 lanes, SST | DHF310478 | US.AVIDIS.Functional.AUXEA-DP-4K | Marc Arends
System Infrastructure | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 6.1 Integration | Supported features | EDID / DDC2, Hot Plug Detect optional | DHF220622 | IDS Appendix C | Marc Arends
System Infrastructure | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 6.1 Integration | Video Output Connection for IP-Based Systems | Tag: US.TechData.VideoOutAdpt |  |  | Marc Arends
System Infrastructure | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 6.1 Integration | Standard | DisplayPort 1.2 | DHF310478 | US.AVIDIS.Functional.AVO-2DP-4K | Marc Arends
System Infrastructure | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 6.1 Integration | Connector type | DisplayPort (connections are limited to a single device) | DHF310478 | US.AVIDIS.Functional.AVO-2DP-4K | Marc Arends
System Infrastructure | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 6.1 Integration | Cable length | 3m DP to DP cable | DHF220622 | IDS.VOIP.Video.Output | Marc Arends
System Infrastructure | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 6.1 Integration | Supported resolutions | 2 x 1920 x 1200 x 60 Hz (FHD) or 1 x 3840 x 2160 x 60 Hz (UHD) | DHF220622 | IDS.VOIP.Video.Output | Marc Arends
System Infrastructure | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 6.1 Integration | Supported clock frequencies | 25 - 600 MHz | DHF220622 | IDS Appendix C | Marc Arends
System Infrastructure | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 6.1 Integration | DP lanes supported | DP 1.2 SST: 2-lanes / DP++ 1.2 SST: 4-lanes (4K) | DHF220622 | IDS.VOIP.Video.Output | Marc Arends
System Infrastructure | - | - | - | - | - | - | - | - | - | - | - | - | 1 | 6.1 Integration | Supported features | DP++ connector supports SL-DVI output | DHF220622 | IDS.VOIP.Video.Output | Marc Arends

## x-ray protocol selection
System | X-ray Protocol |  |  | Ref. AK (mGy/img)
C12/F12 | Pediatric | Pediatrics | 15fps Normal | 0.195
| Cardiac | Left Coronary | 15fps Normal | 0.19636
| Head | Cerebral | 2fps Normal | 5.0511
| Head | Cerebral | 4fps Normal | 5.0494
| Thorax | Lungs | 3fps | 2.522
C12/F12 | Cardio Pediatric | <40 kg Cine | 15 fps low | 0.019
with ClarityIQ (option) | Cardio Pediatric | >40 kg Cine | 15 fps low | 0.037
| Cardiac | Left Coronary | 15fps Low | 0.036218
| Head | Cerebral | 2fps Low | 0.97814
| Head | Cerebral | 4fps Low | 0.97828
| Thorax | Lungs | 3fps | 1.2076
System | X-ray Protocol | Flavor | Ref. AK-rate (mGy/s)
C12/F12 | Pediatric | Low | 0.18264
| Pediatric | Normal | 0.43353
| Pediatric | High | 0.68417
| Cardiac | Low | 0.23761
| Cardiac | Normal | 0.51723
| Cardiac | High | 0.71899
C12/F12 | Cardio Pediatric < 40 kg | Low | 0.06093
with ClarityIQ (option) | Cardio Pediatric < 40 kg | Medium | 0.092038
| Cardio Pediatric < 40 kg | Normal | 0.14401
| Cardiac | Low | 0.1445
| Cardiac | Medium | 0.21351
| Cardiac | Normal | 0.50114
System | X-ray Protocol |  |  | Ref. AK (mGy/img)
C15/F15 | Pediatric | Pediatrics | 15fps Normal | 0.12575
| Cardiac | Left Coronary | 15fps Normal | 0.12466
| Head | Cerebral | 2fps Normal | 2.7849
| Head | Cerebral | 4fps Normal | 2.7843
| Thorax | Lungs | 3fps | 4.312
C15/F15 | Cardio Pediatric | <40 kg Cine | 15 fps low | 0.011659
with ClarityIQ (option) | Cardio Pediatric | >40 kg Cine | 15 fps low | 0.022659
| Cardiac | Left Coronary | 15fps Low | 0.022659
| Head | Cerebral | 2fps Low | 0.57309
| Head | Cerebral | 4fps Low | 0.57248
| Thorax | Lungs | 3fps | 0.70037
System | X-ray Protocol | Flavor | Ref. AK-rate (mGy/s)
C15/F15 | Pediatric | Low | 0.11161
| Pediatric | Normal | 0.24647
| Pediatric | High | 0.41207
| Cardiac | Low | 0.12131
| Cardiac | Normal | 0.35319
| Cardiac | High | 0.42122
C15/F15 | Cardio Pediatric < 40 kg | Low | 0.037741
with ClarityIQ (option) | Cardio Pediatric < 40 kg | Medium | 0.056684
| Cardio Pediatric < 40 kg | Normal | 0.087294
| Cardiac | Low | 0.090185
| Cardiac | Medium | 0.13178
| Cardiac | Normal | 0.3217
System | X-ray Protocol |  |  | Ref. AK (mGy/img)
C20/F20 | Pediatric | Pediatrics | 15fps Normal | 0.10505
| Cardiac | Left Coronary | 15fps Normal | 0.10248
| Head | Cerebral | 2fps Normal | 2.2794
| Head | Cerebral | 4fps Normal | 2.2802
| Thorax | Lungs | 3fps | 1.1403
C20/F20 | Cardio Pediatric | <40 kg Cine | 15 fps low | 0.009665
with ClarityIQ (option) | Cardio Pediatric | >40 kg Cine | 15 fps low | 0.018683
| Cardiac | Left Coronary | 15fps Low | 0.018477
| Head | Cerebral | 2fps Low | 0.46283
| Head | Cerebral | 4fps Low | 0.46269
| Thorax | Subclavian | 3fps | 0.79482
System | X-ray Protocol | Flavor | Ref. AK-rate (mGy/s)
C20/F20 | Pediatric | Low | 0.096027
| Pediatric | Normal | 0.217
| Pediatric | High | 0.38444
| Cardiac | Low | 0.1011
| Cardiac | Normal | 0.29611
| Cardiac | High | 0.3659
C20/F20 | Cardio Pediatric < 40 kg | Low | 0.032459
with ClarityIQ (option) | Cardio Pediatric < 40 kg | Medium | 0.048201
| Cardio Pediatric < 40 kg | Normal | 0.074186
| Cardiac | Low | 0.076229
| Cardiac | Medium | 0.11313
| Cardiac | Normal | 0.27108

## patient type
System | Patient Type | Ref. AK (mGy/img)
C12/F12 | Neonate | 0.040845
* Cardiac procedure | Infant | 0.073555
* Left Coronary | Child, Small Adult | 0.11704
* 15 fps Normal | Default | 0.19636
| Large Adult, Very Large Adult | 0.19664
C12/F12 | Neonate | 0.043273
* Pediatric procedure | Infant | 0.073805
* Pediatric | Child, Small Adult | 0.11548
* 15 fps Normal | Default | 0.19516
System | Patient Type | Ref. AK (mGy/img)
C15/F15 | Neonate | 0.023759
* Cardiac procedure | Infant | 0.045702
* Left Coronary | Child, Small Adult | 0.071729
* 15 fps Normal | Default | 0.12466
| Large Adult, Very Large Adult | 0.11035
C15/F15 | Neonate | 0.024014
* Pediatric procedure | Infant | 0.039191
* Pediatric | Child, Small Adult | 0.072454
* 15 fps Normal | Default | 0.12575
System | Patient Type | Ref. AK (mGy/img)
C20/F20 | Neonate | 0.020295
* Cardiac procedure | Infant | 0.038699
* Left Coronary | Child, Small Adult | 0.059412
* 15 fps Normal | Default | 0.10248
| Large Adult, Very Large Adult | 0.0891
C20/F20 | Neonate | 0.020552
* Pediatric procedure | Infant | 0.0342
* Pediatric | Child, Small Adult | 0.06052
* 15 fps Normal | Default | 0.10505

## field size
C12/F12 With ClarityIQ (option) |  |  | Dose ratio (%) | 100 | 110 | 130 | 155 | 185
|  |  | Field Size (cm) | 30 | 27 | 22 | 19 | 15
X-ray Protocol |  |  | Patient Type | Ref. AK (mGy/img)
head | Cerebral | 2fps Low | Default | 0.97814 | 1.0862 | 1.3035 | 1.5822 | 1.9335
C15/F15 With ClarityIQ (option) |  |  | Dose ratio (%) | 100 | 110 | 130 | 150 | 180 | 215 | 260
|  |  | Field Size (cm) | 39 | 37 | 31 | 27 | 22 | 19 | 15
X-ray Protocol |  |  | Patient Type | Ref. AK (mGy/img)
head | Cerebral | 2fps Low | Default | 0.57309 | 0.63117 | 0.75391 | 0.88483 | 1.0756 | 1.3148 | 1.6248
C20/F20 With ClarityIQ (option) |  |  | Dose ratio (%) | 100 | 130 | 145 | 170 | 200 | 240 | 280 | 330
|  |  | Field Size (cm) | 48 | 42 | 37 | 31 | 27 | 22 | 19 | 15
X-ray Protocol |  |  | Patient Type | Ref. AK (mGy/img)
head | Cerebral | 2fps Low | Default | 0.46283 | 0.60509 | 0.67656 | 0.79819 | 0.94586 | 1.1487 | 1.3582 | 1.6203

## Oblique
|  |  | Patient Thickness | 20 cm PMMA | 30 cm PMMA
System | X-ray Protocol |  |  | Ref. AK (mGy/img) | AK (mGy/img)
C12/F12 | Cardiac | Left Coronary | 15fps Normal | 0.19636 | 0.91168
| Head | Cerebral | 2fps Normal | 5.0511 | 14.0677
| Thorax | Lungs | 3fps | 2.522 | 7.4912
C12/F12 | Cardiac | Left Coronary | 15fps Low | 0.036218 | 0.19956
with ClarityIQ (option) | Head | Cerebral | 2fps Low | 0.97814 | 6.7488
| Thorax | Lungs | 3fps | 1.2076 | 6.4604
|  |  | Patient Thickness | 20 cm PMMA | 30 cm PMMA
System | X-ray Protocol |  |  | Ref. AK (mGy/img) | AK (mGy/img)
C15/F15 | Cardiac | Left Coronary | 15fps Normal | 0.12466 | 0.59969
| Head | Cerebral | 2fps Normal | 2.7849 | 12.9186
| Thorax | Lungs | 3fps | 1.3989 | 5.2085
C15/F15 | Cardiac | Left Coronary | 15fps Low | 0.022659 | 0.12284
with ClarityIQ (option) | Head | Cerebral | 2fps Low | 0.57309 | 4.0177
| Thorax | Lungs | 3fps | 0.70037 | 4.1189
|  |  | Patient Thickness | 20 cm PMMA | 30 cm PMMA
System | X-ray Protocol |  |  | Ref. AK (mGy/img) | AK (mGy/img)
C20/F20 | Cardiac | Left Coronary | 15fps Normal | 0.10248 | 0.55789
| Head | Cerebral | 2fps Normal | 2.2794 | 12.305
| Thorax | Lungs | 3fps | 1.1403 | 4.9564
C20/F20 | Cardiac | Left Coronary | 15fps Low | 0.018477 | 0.11025
with ClarityIQ (option) | Head | Cerebral | 2fps Low | 0.46283 | 3.7648
| Thorax | Lungs | 3fps | 0.56943 | 3.8613

## 2Gy limit
|  | Patient Thickness |  | 20 cm PMMA |  |  | 30 cm PMMA
System | X-ray Protocol |  |  | Ref. AK (mGy/img) | # Exp. Needed | Time at constant fps | Ref. AK (mGy/img) | # Exp. Needed | Time at constant fps
C12/F12 | Cardiac | Left Coronary | 15fps Normal | 0.19636 | 10185.4257 | 11 min | 0.91168 | 2193.7474 | 2.4 min
| Head | Cerebral | 2fps Normal | 5.0511 | 395.9495 | 3.3 min | 14.0677 | 142.1698 | 1.2 min
| Head | Cerebral | 4fps Normal | 5.0494 | 396.0901 | 1.7 min | 15.4134 | 129.7575 | 0.5 min
| Thorax | Lungs | 3fps | 2.522 | 793.0161 | 4.4 min | 7.4912 | 266.9804 | 1.5 min
C12/F12 | Cardiac | Left Coronary | 15fps Low | 0.036218 | 55221.1607 | 61.4 min | 0.19956 | 10021.8978 | 11.1 min
with ClarityIQ (option) | Head | Cerebral | 2fps Low | 0.97814 | 2044.6887 | 17 min | 6.7488 | 296.3503 | 2.5 min
| Head | Cerebral | 4fps Low | 0.97828 | 2044.4003 | 8.5 min | 6.7488 | 296.3503 | 1.2 min
| Thorax | Lungs | 3fps | 1.2076 | 1656.205 | 9.2 min | 6.4604 | 309.5802 | 1.7 min
|  | Patient Thickness | 20 cm PMMA |  | 30 cm PMMA
System | X-ray Protocol | Flavor | Ref. AK-rate (mGy/s) | Time required | Ref. AK-rate (mGy/s) | Time required
C12/F12 | Cardiac | Low | 0.23761 | 140 min | 1.3584 | 25 min
| Cardiac | Normal | 0.51723 | 64 min | 2.2513 | 15 min
| Head | Low | 0.18877 | 177 min | 0.87781 | 38 min
C12/F12 | Cardiac | Low | 0.1445 | 231 min | 0.64423 | 52 min
with ClarityIQ (option) | Cardiac | Normal | 0.50114 | 67 min | 2.2693 | 15 min
| Head | Low | 0.18451 | 181 min | 0.43053 | 77 min
|  | Patient Thickness |  | 20 cm PMMA |  |  | 30 cm PMMA
System | X-ray Protocol |  |  | Ref. AK (mGy/img) | # Exp. Needed | Time at constant fps | Ref. AK (mGy/img) | # Exp. Needed | Time at constant fps
C15/F15 | Cardiac | Left Coronary | 15fps Normal | 0.12466 | 16043.8961 | 17.8 min | 0.59969 | 3335.0398 | 3.7 min
| Head | Cerebral | 2fps Normal | 2.7849 | 718.1574 | 6 min | 9.4779 | 211.0169 | 1.8 min
| Head | Cerebral | 4fps Normal | 2.7843 | 718.307 | 5 min | 9.8429 | 203.193 | 0.8 min
| Thorax | Lungs | 3fps | 4.312 | 463.82 | 2.6 min | 5.2085 | 383.9895 | 2.1 min
C15/F15 | Cardiac | Left Coronary | 15fps Low | 0.022659 | 88265.1485 | 98.1 min | 0.12284 | 16281.3416 | 18.1 min
with ClarityIQ (option) | Head | Cerebral | 2fps Low | 0.57309 | 3489.8715 | 29.1 min | 4.0177 | 497.7992 | 4,1 min
| Head | Cerebral | 4fps Low | 0.57248 | 3493.5535 | 14.6 min | 3.504 | 570.7818 | 2.4 min
| Thorax | Lungs | 3fps | 0.70037 | 2855.6538 | 15.9 min | 4.1189 | 485.5626 | 2.7 min
|  | Patient Thickness | 20 cm PMMA |  | 30 cm PMMA
System | X-ray Protocol | Flavor | Ref. AK-rate (mGy/s) | Time required | Ref. AK-rate (mGy/s) | Time required
C15/F15 | Cardiac | Low | 0.12131 | 275 min | 0.72492 | 46 min
| Cardiac | Normal | 0.35319 | 94 min | 1.4459 | 23 min
| Head | Low | 0.16887 | 197 min | 0.7223 | 46 min
C15/F15 | Cardiac | Low | 0.090185 | 370 min | 0.42146 | 79 min
with ClarityIQ (option) | Cardiac | Normal | 0.3217 | 104 min | 1.5249 | 22 min
| Head | Low | 0.12118 | 275 min | 0.36238 | 92 min
|  | Patient Thickness |  | 20 cm PMMA |  |  | 30 cm PMMA
System | X-ray Protocol |  |  | Ref. AK (mGy/img) | # Exp. Needed | Time at constant fps | Ref. AK (mGy/img) | # Exp. Needed | Time at constant fps
C20/F20 | Cardiac | Left Coronary | 15fps Normal | 0.10248 | 19515.2414 | 21.7 min | 0.55789 | 3584.9297 | 4.0 min
| Head | Cerebral | 2fps Normal | 2.2794 | 877.4385 | 7.3 min | 12.3046 | 162.5409 | 1.4 min
| Head | Cerebral | 4fps Normal | 2.2802 | 877.1249 | 3.7 min | 10.0644 | 198.7208 | 0,8 min
| Thorax | Lungs | 3fps | 1.1403 | 1753.9998 | 9.7 min | 4.9564 | 403.5153 | 2.2 min
C20/F20 | Cardiac | Left Coronary | 15fps Low | 0.018477 | 108242.6801 | 120.3 min | 0.11025 | 18140.9187 | 20.2 min
with ClarityIQ (option) | Head | Cerebral | 2fps Low | 0.46283 | 4321.2411 | 36 min | 3.7648 | 531.2301 | 4.4 min
| Head | Cerebral | 4fps Low | 0.46269 | 4322.5299 | 18 min | 3.2832 | 609.1616 | 2.5 min
| Thorax | Lungs | 3fps | 0.56943 | 3512.2657 | 19.5 min | 3.8613 | 517.9591 | 2.9 min
|  | Patient Thickness | 20 cm PMMA |  | 30 cm PMMA
System | X-ray Protocol | Flavor | Ref. AK-rate (mGy/s) | Time required | Ref. AK-rate (mGy/s) | Time required
C20/F20 | Cardiac | Low | 0.1011 | 330 min | 0.6564 | 51 min
| Cardiac | normal | 0.29611 | 113 min | 1.328 | 25 min
| Head | Low | 0.3659 | 91 min | 0.60063 | 55 min
C20/F20 | Cardiac | Low | 0.076229 | 437 min | 0.38732 | 86 min
with ClarityIQ (option) | Cardiac | normal | 0.11313 | 295 min | 1.415 | 24 min
| Head | Low | 0.27108 | 123 min | 0.36263 | 92 min

## FD12
ch 16.20.1 | C12/F12 Systems
ch 16.20.1.1 | RAK values | 200 mm PMMA | Large SID
| Radioscopy | (Fluoro) | All modes for all application groups
|  | Field of View | 0 | 1 | 2 | 3 | 4
|  | Field Size (cm) | 30 | 27 | 22 | 19 | 15
| Flavor | Patient Type | Ref. AK (mGy/s)
Cardio | Low | Default | 0.23301 | 0.25446 | 0.2963 | 0.34946 | 0.42012
| Norm | Default | 0.51156 | 0.66403 | 0.84957 | 0.86087 | 0.87621
| High | Default | 0.70902 | 0.77016 | 0.88446 | 1.0258 | 1.1829
Pediatrics | Low | Default | 0.18264 | 0.19939 | 0.23308 | 0.27555 | 0.32833
| Norm | Default | 0.43298 | 0.46974 | 0.53864 | 0.61075 | 0.69784
| High | Default | 0.68428 | 0.73105 | 0.79223 | 0.86437 | 0.95107
EP | Low | Default | 0.086842 | 0.095231 | 0.11216 | 0.13186 | 0.15659
| Norm | Default | 0.20609 | 0.22448 | 0.26146 | 0.30813 | 0.36535
| High | Default | 0.24697 | 0.26934 | 0.31304 | 0.36813 | 0.43639
EP Mapping | Low | Default | 0.043419 | 0.047711 | 0.056487 | 0.067544 | 0.081573
| Norm | Default | 0.10293 | 0.11224 | 0.13069 | 0.15436 | 0.18405
| High | Default | 0.12348 | 0.13478 | 0.15659 | 0.18394 | 0.21822
Head | Low | Default | 0.18425 | 0.25211 | 0.29081 | 0.29939 | 0.31022
| Norm | Default | 0.47158 | 0.50321 | 0.56472 | 0.64077 | 0.73399
| High | Default | 0.75284 | 0.81494 | 0.93732 | 1.0941 | 1.2778
Spine | Low | Default | 0.18425 | 0.25211 | 0.29081 | 0.29939 | 0.31022
| Norm | Default | 0.47158 | 0.50321 | 0.56472 | 0.64077 | 0.73399
| High | Default | 0.75284 | 0.81494 | 0.93732 | 1.0941 | 1.2778
Thorax | Low | Default | 0.18425 | 0.25211 | 0.29081 | 0.29939 | 0.31022
| Norm | Default | 0.47158 | 0.50321 | 0.56472 | 0.64077 | 0.73399
| High | Default | 0.75284 | 0.81494 | 0.93732 | 1.0941 | 1.2778
Abdomen | Low | Default | 0.18425 | 0.25211 | 0.29081 | 0.29939 | 0.31022
| Norm | Default | 0.47158 | 0.50321 | 0.56472 | 0.64077 | 0.73399
| High | Default | 0.75284 | 0.81494 | 0.93732 | 1.0941 | 1.2778
Peripheral | Low | Default | 0.18425 | 0.25211 | 0.29081 | 0.29939 | 0.31022
| Norm | Default | 0.47158 | 0.50321 | 0.56472 | 0.64077 | 0.73399
| High | Default | 0.75284 | 0.81494 | 0.93732 | 1.0941 | 1.2778
Radiography | Exposure | typical and worst case samples per application group
|  | field of View | 0 | 1 | 2 | 3 | 4
|  | Field Size (cm): | 30 | 27 | 22 | 19 | 15
| X-ray Protocol | Patient Type | Ref. AK (mGy/image)
Cardio | Left Coronary | Default | 0.10377 | 0.11258 | 0.13029 | 0.15187 | 0.17787
| 15 fps Low | Infant | 0.073805 | 0.078544 | 0.087842 | 0.099428 | 0.11353
|  | Large Adult, VeryLarge Adult | 0.10369 | 0.11263 | 0.13037 | 0.15192 | 0.17797
|  | Neonate | 0.040121 | 0.043819 | 0.051496 | 0.061546 | 0.074395
| Rotational Scan | Default | 0.19519 | 0.21181 | 0.24373 | 0.28212 | 0.32934
| Prop Ang0 4s | Infant | 0.063269 | 0.068207 | 0.077879 | 0.090131 | 0.10568
|  | Neonate | 0.040121 | 0.043819 | 0.051496 | 0.061546 | 0.074395
Pediatrics | 15 fps Contrast Normal | Default | 0.19516 | 0.21185 | 0.24377 | 0.28211 | 0.32928
|  | Child,SmallAdult | 0.11548 | 0.12521 | 0.1449 | 0.16912 | 0.19809
|  | Infant | 0.073805 | 0.078544 | 0.087842 | 0.099428 | 0.11353
|  | Neonate | 0.043334 | 0.046926 | 0.054155 | 0.063591 | 0.075924
| Prop Free Position | Default | 0.20113 | 0.21835 | 0.25218 | 0.29298 | 0.34168
|  | Neonate,Infant | 0.043334 | 0.046926 | 0.054155 | 0.063591 | 0.075924
EP | 15 Fps | Default | 0.11548 | 0.12521 | 0.1449 | 0.16912 | 0.19809
| Prop 4s | Default | 0.22823 | 0.23056 | 0.2346 | 0.23976 | 0.24627
|  | Neonate,Infant | 0.14768 | 0.14955 | 0.15242 | 0.1563 | 0.16075
|  | LargeAdult,VeryLargeAdult | 0.22823 | 0.23056 | 0.2346 | 0.23976 | 0.24627
EP Mapping | 7.5 fps low | Default | 0.016591 | 0.018228 | 0.021467 | 0.025608 | 0.030723
| 3D EP Prop 4s | Default | 0.22823 | 0.23056 | 0.2346 | 0.23976 | 0.24627
|  | Neonate,Infant | 0.14768 | 0.14955 | 0.15242 | 0.1563 | 0.16075
|  | LargeAdult,VeryLargeAdult | 0.22823 | 0.23056 | 0.2346 | 0.23976 | 0.24627
Head | Cerebral 2fps normal | Default | 4.895 | 5.436 | 6.5402 | 7.7015 | 8.5024
|  | Child | 3.1525 | 3.3776 | 3.8201 | 4.3787 | 5.0724
|  | Neonate,Infant | 3.1525 | 3.3776 | 3.8201 | 4.3787 | 5.0724
| Aortic Arch 3 fps | Default | 7.3036 | 8.1135 | 9.7134 | 11.7876 | 12.8945
Spine | 4 fps | Default | 2.0496 | 2.2788 | 2.7391 | 3.3353 | 4.0877
| 2 fps | Default | 2.5655 | 2.8512 | 3.4278 | 4.1721 | 5.1126
Thorax | Lungs 2 fps | Default | 2.446 | 2.7185 | 3.2691 | 3.9505 | 4.361
| Subclavian 3 fps | Default | 9.046 | 10.6732 | 10.8101 | 10.9715 | 11.1828
Abdomen | Abdomen | Default | 3.8242 | 4.2416 | 5.1008 | 6.1983 | 7.5784
| 3 fps | LargeAdult, VeryLargeAdult | 3.3166 | 3.681 | 4.4337 | 5.3978 | 6.6204
| Iliac / Pelvis 3fps | Default | 7.2974 | 8.1082 | 9.7297 | 11.8418 | 14.4718
Peripheral | Upper Legs 3 fps | Default | 6.6081 | 7.3341 | 8.209 | 8.9328 | 9.7813
| Lower Legs 1 fps | Default | 8.2287 | 8.6443 | 9.3197 | 10.1247 | 11.0434
Roadmap |  | All modes for all application groups
|  | field of View | 0 | 1 | 2 | 3 | 4
|  | Field Size (cm): | 30 | 27 | 22 | 19 | 15
| mode | Patient Type | Ref. AK (mGy/image)
Cardio | Navigate | Default | 0.38281 | 0.48589 | 0.49039 | 0.49568 | 0.50257
| UnSubtract | Default | 0.38281 | 0.48589 | 0.49039 | 0.49568 | 0.50257
Pediatrics | Navigate | Default | 0.38281 | 0.48589 | 0.49039 | 0.49568 | 0.50257
| UnSubtract | Default | 0.38281 | 0.48589 | 0.49039 | 0.49568 | 0.50257
EP | Navigate | Default | 0.38281 | 0.48589 | 0.49039 | 0.49568 | 0.50257
EP Mapping | Navigate | Default | 0.38281 | 0.48589 | 0.49039 | 0.49568 | 0.50257
Head | Navigate | Default | 0.38281 | 0.48589 | 0.49039 | 0.49568 | 0.50257
| Carotid | Default | 0.38058 | 0.46141 | 0.46528 | 0.47036 | 0.47561
| Coil | Default | 1.8177 | 2.0195 | 2.4279 | 2.9511 | 3.6108
Spine | UnSubtract | Default | 0.38281 | 0.48589 | 0.49039 | 0.49568 | 0.50257
| Navigate | Default | 0.38281 | 0.48589 | 0.49039 | 0.49568 | 0.50257
| Stent | Default | 0.38058 | 0.46141 | 0.46528 | 0.47036 | 0.47561
Thorax | UnSubtract | Default | 0.38281 | 0.48589 | 0.49039 | 0.49568 | 0.50257
| Navigate | Default | 0.38281 | 0.48589 | 0.49039 | 0.49568 | 0.50257
| Intervention | Default | 0.38058 | 0.46141 | 0.46528 | 0.47036 | 0.47561
Abdomen | UnSubtract | Default | 0.38281 | 0.48589 | 0.49039 | 0.49568 | 0.50257
| Navigate | Default | 0.38281 | 0.48589 | 0.49039 | 0.49568 | 0.50257
| Stent | Default | 0.38058 | 0.46141 | 0.46528 | 0.47036 | 0.47561
Peripheral | UnSubtract | Default | 0.38281 | 0.48589 | 0.49039 | 0.49568 | 0.50257
| Navigate | Default | 0.38281 | 0.48589 | 0.49039 | 0.49568 | 0.50257

## FD12 clarity
ch 16.20.2 | C12/F12 Systems | Clarity
ch 16.20.2.1 | RAK values | 200 mm PMMA | Large SID
| Radioscopy | (Fluoro) | All modes for all application groups
|  | Field of View | 0 | 1 | 2 | 3 | 4
|  | Field Size (cm) | 30 | 27 | 22 | 19 | 15
| Flavor | Patient Type | Ref. AK (mGy/s)
Cardio Clarity | Low | Default | 0.14706 | 0.15932 | 0.184 | 0.21366 | 0.24972
| Medium | Default | 0.21664 | 0.23598 | 0.27409 | 0.32157 | 0.38007
| Normal | Default | 0.51179 | 0.55279 | 0.63488 | 0.7371 | 0.85846
Cardio Pediatrics Clarity <40 kg | Low | Default | 0.06093 | 0.066483 | 0.077564 | 0.091439 | 0.10864
| Medium | Default | 0.092038 | 0.10026 | 0.11591 | 0.13527 | 0.15771
| Normal | Default | 0.14401 | 0.15677 | 0.18098 | 0.21103 | 0.24769
Cardio Pediatrics Clarity >40 kg | Low | Default | 0.14706 | 0.15932 | 0.184 | 0.21366 | 0.24972
| Medium | Default | 0.21664 | 0.23598 | 0.27409 | 0.32157 | 0.38007
| Normal | Default | 0.51179 | 0.55279 | 0.63488 | 0.7371 | 0.85846
EP Clarity | Low | Default | 0.043606 | 0.047359 | 0.054727 | 0.063915 | 0.072768
| Medium | Default | 0.086443 | 0.093945 | 0.10857 | 0.12678 | 0.14713
| Normal | Default | 0.10613 | 0.11606 | 0.13503 | 0.15906 | 0.18866
EP Mapping Clarity | Low | Default | 0.043606 | 0.047359 | 0.054727 | 0.063915 | 0.072768
| Medium | Default | 0.086443 | 0.093945 | 0.10857 | 0.12678 | 0.14713
| Normal | Default | 0.10613 | 0.11606 | 0.13503 | 0.15906 | 0.18866
Head Clarity | Low | Default | 0.1886 | 0.20226 | 0.22932 | 0.26343 | 0.30709
| Medium | Default | 0.30582 | 0.3337 | 0.38933 | 0.44716 | 0.51394
| Normal | Default | 0.55654 | 0.60362 | 0.69734 | 0.81264 | 0.95099
Spine Clarity | Low | Default | 0.1886 | 0.20226 | 0.22932 | 0.26343 | 0.30709
| Medium | Default | 0.30582 | 0.3337 | 0.38933 | 0.44716 | 0.51394
| Normal | Default | 0.55654 | 0.60362 | 0.69734 | 0.81264 | 0.95099
Thorax Clarity | Low | Default | 0.1886 | 0.20226 | 0.22932 | 0.26343 | 0.30709
| Medium | Default | 0.30582 | 0.3337 | 0.38933 | 0.44716 | 0.51394
| Normal | Default | 0.55654 | 0.60362 | 0.69734 | 0.81264 | 0.95099
Abdomen Clarity | Low | Default | 0.1886 | 0.20226 | 0.22932 | 0.26343 | 0.30709
| Medium | Default | 0.30582 | 0.3337 | 0.38933 | 0.44716 | 0.51394
| Normal | Default | 0.55654 | 0.60362 | 0.69734 | 0.81264 | 0.95099
Peripheral Clarity | Low | Default | 0.14706 | 0.15932 | 0.184 | 0.21366 | 0.24972
| Medium | Default | 0.21664 | 0.23598 | 0.27409 | 0.32157 | 0.38007
| Normal | Default | 0.51179 | 0.55279 | 0.63488 | 0.7371 | 0.85846
Radiography | Exposure | typical and worst case samples per application group
|  | field of View | 0 | 1 | 2 | 3 | 4
|  | Field Size (cm): | 30 | 27 | 22 | 19 | 15
| X-ray Protocol | Patient Type | Ref. AK (mGy/image)
Cardio Clarity | Left Coronary 15 fps Low | Default | 0.036635 | 0.040075 | 0.046662 | 0.054835 | 0.064879
| Rotational Scan Prop Ang0 4s | Default | 0.19519 | 0.21181 | 0.24373 | 0.28212 | 0.32934
|  | Infant | 0.063269 | 0.068207 | 0.077879 | 0.090131 | 0.10568
|  | Neonate | 0.040121 | 0.043819 | 0.051496 | 0.061546 | 0.074395
Cardio Pediatrics Clarity <40 kg | 15 fps Contrast low | Default | 0.01854 | 0.020178 | 0.023419 | 0.027562 | 0.032418
| 15 fps Contrast normal | Default | 0.030122 | 0.033021 | 0.038694 | 0.045989 | 0.054752
Cardio Pediatrics Clarity >40 kg | 15 fps Contrast low | Default | 0.036635 | 0.040075 | 0.046662 | 0.054835 | 0.064879
|  | Neonate,Infant,Child | 0.01854 | 0.020178 | 0.023419 | 0.027562 | 0.032418
| 15 fps Contrast normal | Default | 0.077016 | 0.083838 | 0.096932 | 0.11321 | 0.13281
|  | Neonate,Infant,Child | 0.030122 | 0.033021 | 0.038694 | 0.045989 | 0.054752
EP Clarity | 15 Fps | Default | 0.036635 | 0.040075 | 0.046662 | 0.054835 | 0.064879
| Prop 4s | Default | 0.22823 | 0.23056 | 0.2346 | 0.23976 | 0.24627
|  | Neonate,Infant | 0.14768 | 0.14955 | 0.15242 | 0.1563 | 0.16075
|  | LargeAdult,VeryLargeAdult | 0.22823 | 0.23056 | 0.2346 | 0.23976 | 0.24627
EP Mapping Clarity | 7.5 fps | Default | 0.036635 | 0.040075 | 0.046662 | 0.054835 | 0.064879
| 3D EP Prop 4s | Default | 0.22823 | 0.23056 | 0.2346 | 0.23976 | 0.24627
|  | Neonate,Infant | 0.14768 | 0.14955 | 0.15242 | 0.1563 | 0.16075
|  | LargeAdult,VeryLargeAdult | 0.22823 | 0.23056 | 0.2346 | 0.23976 | 0.24627
Head Clarity | Cerebral 2fps low | Default | 0.99034 | 1.0986 | 1.3183 | 1.6019 | 1.9533
| Aortic Arch 3 fps | Default | 1.4227 | 1.5797 | 1.8953 | 2.3051 | 2.8099
Spine Clarity | 4 fps | Default | 1.1326 | 1.26 | 1.5194 | 1.8523 | 2.2728
| 2 fps | Default | 1.1335 | 1.2601 | 1.5182 | 1.8526 | 2.2725
Thorax Clarity | Lungs 2 fps | Default | 1.2209 | 1.3574 | 1.6298 | 1.9856 | 2.4316
| Subclavian 3 fps | Default | 1.7217 | 1.9115 | 2.2886 | 2.7736 | 3.3778
Abdomen Clarity | 3 fps low | Default | 0.79655 | 0.88435 | 1.0596 | 1.2853 | 1.5686
| Iliac / Pelvis 3fps | Default | 3.8904 | 4.3221 | 5.2019 | 6.3402 | 7.1376
Peripheral Clarity | Upper Legs 3 fps | Default | 1.128 | 1.2535 | 1.5015 | 1.6883 | 1.9032
| Lower Legs 1 fps | Default | 2.0635 | 2.2854 | 2.7318 | 3.3018 | 4.0056
Roadmap |  | All modes for all application groups
|  | field of View | 0 | 1 | 2 | 3 | 4
|  | Field Size (cm): | 30 | 27 | 22 | 19 | 15
| mode | Patient Type | Ref. AK (mGy/image)
Cardio Clarity | Navigate | Default | 0.17516 | 0.18801 | 0.21264 | 0.24323 | 0.27906
| UnSubtract | Default | 0.17516 | 0.18801 | 0.21264 | 0.24323 | 0.27906
Cardio Pediatrics Clarity <40 kg | Navigate | Default | 0.17516 | 0.18801 | 0.21264 | 0.24323 | 0.27906
| UnSubtract | Default | 0.17516 | 0.18801 | 0.21264 | 0.24323 | 0.27906
Cardio Pediatrics Clarity >40 kg | Navigate | Default | 0.17516 | 0.18801 | 0.21264 | 0.24323 | 0.27906
| UnSubtract | Default | 0.17516 | 0.18801 | 0.21264 | 0.24323 | 0.27906
EP Clarity | Navigate | Default | 0.17516 | 0.18801 | 0.21264 | 0.24323 | 0.27906
| UnSubtract | Default | 0.17516 | 0.18801 | 0.21264 | 0.24323 | 0.27906
EP Mapping Clarity | Navigate | Default | 0.17516 | 0.18801 | 0.21264 | 0.24323 | 0.27906
| UnSubtract | Default | 0.17516 | 0.18801 | 0.21264 | 0.24323 | 0.27906
Head Clarity | Navigate | Default | 0.17516 | 0.18801 | 0.21264 | 0.24323 | 0.27906
| Carotid | Default | 0.33483 | 0.37173 | 0.44659 | 0.54229 | 0.6613
| Coil | Default | 0.17516 | 0.18801 | 0.21264 | 0.24323 | 0.27906
Spine Clarity | UnSubtract | Default | 0.17516 | 0.18801 | 0.21264 | 0.24323 | 0.27906
| Navigate | Default | 0.2361 | 0.26218 | 0.31428 | 0.38168 | 0.46659
| Stent | Default | 0.17516 | 0.18801 | 0.21264 | 0.24323 | 0.27906
Thorax Clarity | UnSubtract | Default | 0.17516 | 0.18801 | 0.21264 | 0.24323 | 0.27906
| Navigate | Default | 0.17516 | 0.18801 | 0.21264 | 0.24323 | 0.27906
Abdomen Clarity | UnSubtract | Default | 0.17516 | 0.18801 | 0.21264 | 0.24323 | 0.27906
| Navigate | Default | 0.17516 | 0.18801 | 0.21264 | 0.24323 | 0.27906
Peripheral Clarity | UnSubtract | Default | 0.17516 | 0.18801 | 0.21264 | 0.24323 | 0.27906
| Navigate | Default | 0.17516 | 0.18801 | 0.21264 | 0.24323 | 0.27906

## FD15
ch 16.20.3 | F15 Systems
ch 16.20.3.1 | RAK values | 200 mm PMMA | Large SID
| Radioscopy | (Fluoro) | All modes for all application groups
|  | Field of View | 0 | 1 | 2 | 3 | 4 | 5 | 6
|  | Field Size (cm) | 39 | 37 | 31 | 27 | 22 | 19 | 15
| Flavor | Patient Type | All modes for all application groups 0: Ref. AK (mGy/s)
Cardio | Low | Default | 0.12131 | 0.13222 | 0.15502 | 0.17875 | 0.21363 | 0.25146 | 0.31272
| Norm | Default | 0.35319 | 0.37989 | 0.43188 | 0.4843 | 0.55702 | 0.6342 | 0.75436
| High | Default | 0.42122 | 0.45365 | 0.51877 | 0.58642 | 0.68035 | 0.78006 | 0.93586
Pediatrics | Low | Default | 0.11157 | 0.12081 | 0.1403 | 0.16057 | 0.19009 | 0.227 | 0.27404
| Norm | Default | 0.24647 | 0.26604 | 0.30677 | 0.34876 | 0.4031 | 0.46578 | 0.54582
| High | Default | 0.41169 | 0.43985 | 0.49418 | 0.54938 | 0.62622 | 0.71153 | 0.78561
EP | Low | Default | 0.050518 | 0.055036 | 0.064625 | 0.074874 | 0.089661 | 0.10781 | 0.12982
| Norm | Default | 0.12083 | 0.13155 | 0.15371 | 0.17562 | 0.20729 | 0.24773 | 0.3002
| High | Default | 0.14367 | 0.15661 | 0.18359 | 0.21199 | 0.25324 | 0.30494 | 0.37163
EP Mapping | Low | Default | 0.025245 | 0.027515 | 0.032361 | 0.037452 | 0.044824 | 0.054216 | 0.066403
| Norm | Default | 0.060418 | 0.065755 | 0.07685 | 0.087844 | 0.10371 | 0.12382 | 0.14997
| High | Default | 0.071857 | 0.078275 | 0.091928 | 0.10611 | 0.12663 | 0.1524 | 0.18575
Head | Low | Default | 0.16887 | 0.18309 | 0.2135 | 0.24598 | 0.2933 | 0.35477 | 0.43518
| Norm | Default | 0.28658 | 0.30918 | 0.35589 | 0.40344 | 0.47256 | 0.55041 | 0.63975
| High | Default | 0.38773 | 0.41696 | 0.47669 | 0.53712 | 0.62367 | 0.72809 | 0.85916
Spine | Low | Default | 0.13197 | 0.14286 | 0.16627 | 0.19104 | 0.22706 | 0.2732 | 0.33365
| Norm | Default | 0.28496 | 0.30736 | 0.35381 | 0.40143 | 0.46982 | 0.54739 | 0.63614
| High | Default | 0.47331 | 0.50841 | 0.58308 | 0.65701 | 0.76334 | 0.89242 | 1.0541
Thorax | Low | Default | 0.1232 | 0.13343 | 0.1553 | 0.17826 | 0.21137 | 0.25427 | 0.31001
| Norm | Default | 0.28658 | 0.30918 | 0.35589 | 0.40344 | 0.47256 | 0.55041 | 0.63975
| High | Default | 0.4758 | 0.51184 | 0.58668 | 0.661 | 0.76834 | 0.89736 | 1.0601
Abdomen | Low | Default | 0.13197 | 0.14286 | 0.16627 | 0.19104 | 0.22706 | 0.2732 | 0.33365
| Norm | Default | 0.28496 | 0.30736 | 0.35381 | 0.40143 | 0.46982 | 0.54739 | 0.63614
| High | Default | 0.47331 | 0.50841 | 0.58308 | 0.65701 | 0.76334 | 0.89242 | 1.0541
Peripheral | Low | Default | 0.13197 | 0.14286 | 0.16627 | 0.19104 | 0.22706 | 0.2732 | 0.33365
| Norm | Default | 0.28496 | 0.30736 | 0.35381 | 0.40143 | 0.46982 | 0.54739 | 0.63614
| High | Default | 0.47331 | 0.50841 | 0.58308 | 0.65701 | 0.76334 | 0.89242 | 1.0541
Radiography | Exposure | typical and worst case samples per application group
|  | field of View | 0 | 1 | 2 | 3 | 4 | 5 | 6
|  | Field Size (cm): | 39 | 37 | 31 | 27 | 22 | 19 | 15
| X-ray Protocol | Patient Type | 0: Ref. AK (mGy/image)
Cardio | Left Coronary | Default | 0.063893 | 0.069221 | 0.07981 | 0.090564 | 0.10627 | 0.1245 | 0.14755
| 15 fps Low | Infant | 0.045702 | 0.049183 | 0.054648 | 0.060193 | 0.067973 | 0.077634 | 0.089882
|  | Large Adult, VeryLarge Adult | 0.063179 | 0.068395 | 0.078885 | 0.089901 | 0.10513 | 0.12318 | 0.14609
|  | Neonate | 0.023759 | 0.025702 | 0.029859 | 0.034238 | 0.040878 | 0.049592 | 0.061013
| Rotational Scan | Default | 0.12261 | 0.13243 | 0.15253 | 0.17298 | 0.20133 | 0.23511 | 0.27793
| Prop Ang0 4s | Infant | 0.038864 | 0.041554 | 0.047067 | 0.052862 | 0.061037 | 0.071417 | 0.085174
|  | Neonate | 0.023759 | 0.025702 | 0.029859 | 0.034238 | 0.040878 | 0.049592 | 0.061013
Pediatrics | 15 fps Contrast Normal | Default | 0.12575 | 0.13613 | 0.1565 | 0.17654 | 0.20577 | 0.24052 | 0.28207
|  | Child,SmallAdult | 0.072454 | 0.078447 | 0.090627 | 0.10251 | 0.12015 | 0.14098 | 0.16678
|  | Infant | 0.03919 | 0.041877 | 0.047495 | 0.053273 | 0.061565 | 0.07212 | 0.085917
|  | Neonate | 0.023998 | 0.025954 | 0.030129 | 0.034621 | 0.04132 | 0.050052 | 0.061649
| Prop Free Position | Default | 0.12161 | 0.13123 | 0.1512 | 0.17159 | 0.20086 | 0.23484 | 0.27621
|  | Neonate,Infant | 0.026598 | 0.028504 | 0.032544 | 0.036707 | 0.042719 | 0.050361 | 0.060466
EP | 15 Fps | Default | 0.072454 | 0.078447 | 0.090627 | 0.10251 | 0.12015 | 0.14098 | 0.16678
| 15 Fps | Default | 0.072454 | 0.078447 | 0.090627 | 0.10251 | 0.12015 | 0.14098 | 0.16678
EP Mapping | 7.5 fps low | Default | 0.010059 | 0.010919 | 0.012739 | 0.014615 | 0.017384 | 0.020842 | 0.025223
| 15 fps normal | Default | 0.072469 | 0.078432 | 0.090621 | 0.1025 | 0.12018 | 0.14091 | 0.1669
Head | Cerebral 2fps normal | Default | 2.7849 | 3.0676 | 3.6753 | 4.316 | 5.2597 | 6.448 | 7.9019
|  | Child | 1.7696 | 1.8833 | 2.1243 | 2.3734 | 2.7297 | 3.1983 | 3.838
|  | Neonate,Infant | 1.7696 | 1.8833 | 2.1243 | 2.3734 | 2.7297 | 3.1983 | 3.838
| Aortic Arch 3 fps | Default | 4.312 | 4.7497 | 5.6706 | 6.6351 | 6.7294 | 6.8801 | 7.0201
Spine | 4 fps | Default | 1.1718 | 1.2907 | 1.5476 | 1.8152 | 2.2136 | 2.7141 | 3.3654
| 2 fps | Default | 1.4664 | 1.6153 | 1.934 | 2.2705 | 2.7668 | 3.394 | 4.2092
Thorax | Lungs 2 fps | Default | 1.399 | 1.5426 | 1.8476 | 2.1675 | 2.624 | 2.8956 | 3.234
| Subclavian 3 fps | Default | 5.3379 | 5.881 | 7.014 | 8.2011 | 9.9398 | 11.2542 | 11.3157
Abdomen | Abdomen | Default | 2.6159 | 2.883 | 3.4464 | 4.0472 | 4.9252 | 6.0296 | 6.5536
| 3 fps | LargeAdult, VeryLargeAdult | 2.2642 | 2.4997 | 2.9909 | 3.5199 | 4.2904 | 5.2618 | 6.5381
| Iliac / Pelvis 3fps | Default | 4.2501 | 4.6869 | 5.6053 | 6.5778 | 7.9993 | 9.8053 | 12.1389
Peripheral | Upper Legs 3 fps | Default | 3.7841 | 4.1694 | 4.9735 | 5.5513 | 6.044 | 6.6211 | 7.3262
| Lower Legs 1 fps | Default | 4.7323 | 5.2088 | 5.7742 | 6.1934 | 6.7365 | 7.3435 | 8.0978
Roadmap |  | All modes for all application groups
|  | field of View | 0 | 1 | 2 | 3 | 4 | 5 | 6
|  | Field Size (cm): | 39 | 37 | 31 | 27 | 22 | 19 | 15
| mode | Patient Type | 0: Ref. AK (mGy/image)
Cardio | Navigate | Default | 0.19816 | 0.24368 | 0.26858 | 0.2956 | 0.35055 | 0.41079 | 0.44342
| UnSubtract | Default | 0.19816 | 0.24368 | 0.26858 | 0.2956 | 0.35055 | 0.41079 | 0.44342
Pediatrics | Navigate | Default | 0.19816 | 0.24368 | 0.26858 | 0.2956 | 0.35055 | 0.41079 | 0.44342
| UnSubtract | Default | 0.19816 | 0.24368 | 0.26858 | 0.2956 | 0.35055 | 0.41079 | 0.44342
EP | Navigate | Default | 0.19816 | 0.24368 | 0.26858 | 0.2956 | 0.35055 | 0.41079 | 0.44342
EP Mapping | Navigate | Default | 0.19816 | 0.24368 | 0.26858 | 0.2956 | 0.35055 | 0.41079 | 0.44342
Head | Navigate | Default | 0.2465 | 0.26624 | 0.30651 | 0.33012 | 0.40142 | 0.43738 | 0.44324
| Carotid | Default | 0.2116 | 0.22515 | 0.25247 | 0.27912 | 0.31754 | 0.34177 | 0.34705
| Coil | Default | 1.0642 | 1.1713 | 1.4012 | 1.6452 | 2.0014 | 2.4498 | 3.0367
Spine | UnSubtract | Default | 0.2465 | 0.26624 | 0.30651 | 0.33012 | 0.40142 | 0.43738 | 0.44324
| Navigate | Default | 0.2465 | 0.26624 | 0.30651 | 0.33012 | 0.40142 | 0.43738 | 0.44324
| Stent | Default | 0.2116 | 0.22515 | 0.25247 | 0.27912 | 0.31754 | 0.34177 | 0.34705
Thorax | UnSubtract | Default | 0.2465 | 0.26624 | 0.30651 | 0.33012 | 0.40142 | 0.43738 | 0.44324
| Navigate | Default | 0.2465 | 0.26624 | 0.30651 | 0.33012 | 0.40142 | 0.43738 | 0.44324
| Intervention | Default | 0.2116 | 0.22515 | 0.25247 | 0.27912 | 0.31754 | 0.34177 | 0.34705
Abdomen | UnSubtract | Default | 0.2465 | 0.26624 | 0.30651 | 0.33012 | 0.40142 | 0.43738 | 0.44324
| Navigate | Default | 0.2465 | 0.26624 | 0.30651 | 0.33012 | 0.40142 | 0.43738 | 0.44324
| Stent | Default | 0.2116 | 0.22515 | 0.25247 | 0.27912 | 0.31754 | 0.34177 | 0.34705
Peripheral | UnSubtract | Default | 0.2465 | 0.26624 | 0.30651 | 0.33012 | 0.40142 | 0.43738 | 0.44324
| Navigate | Default | 0.2465 | 0.26624 | 0.30651 | 0.33012 | 0.40142 | 0.43738 | 0.44324

## FD15 Clarity
ch 16.20.4 | F15 Systems | Clarity
ch 16.20.4.1 | RAK values | 200 mm PMMA | Large SID
| Radioscopy | (Fluoro) | All modes for all application groups
|  | Field of View | 0 | 1 | 2 | 3 | 4 | 5 | 6
|  | Field Size (cm) | 39 | 37 | 31 | 27 | 22 | 19 | 15
| Flavor | Patient Type | All modes for all application groups 0: Ref. AK (mGy/s)
Cardio Clarity | Low | Default | 0.090185 | 0.097607 | 0.11246 | 0.12729 | 0.14873 | 0.17472 | 0.20576
| Medium | Default | 0.13178 | 0.14246 | 0.16456 | 0.1869 | 0.21962 | 0.25926 | 0.31027
| Normal | Default | 0.3217 | 0.3462 | 0.39679 | 0.44665 | 0.51734 | 0.60362 | 0.7102
Cardio Pediatrics Clarity <40 kg | Low | Default | 0.037741 | 0.040805 | 0.047111 | 0.053663 | 0.063174 | 0.075097 | 0.09032
| Medium | Default | 0.056702 | 0.061399 | 0.071162 | 0.081233 | 0.095782 | 0.11359 | 0.13648
| Normal | Default | 0.087294 | 0.094495 | 0.10924 | 0.12455 | 0.14631 | 0.17213 | 0.2039
Cardio Pediatrics Clarity >40 kg | Low | Default | 0.090185 | 0.097607 | 0.11246 | 0.12729 | 0.14873 | 0.17472 | 0.20576
| Medium | Default | 0.13178 | 0.14246 | 0.16456 | 0.1869 | 0.21962 | 0.25926 | 0.31027
| Normal | Default | 0.3217 | 0.3462 | 0.39679 | 0.44665 | 0.51734 | 0.60362 | 0.7102
EP Clarity | Low | Default | 0.026717 | 0.028828 | 0.033173 | 0.037607 | 0.043903 | 0.051666 | 0.06092
| Medium | Default | 0.053051 | 0.057244 | 0.065868 | 0.07461 | 0.087194 | 0.10263 | 0.12083
| Normal | Default | 0.063429 | 0.068744 | 0.079952 | 0.091303 | 0.10753 | 0.12779 | 0.15343
EP Mapping Clarity | Low | Default | 0.026717 | 0.028828 | 0.033173 | 0.037607 | 0.043903 | 0.051666 | 0.06092
| Medium | Default | 0.053051 | 0.057244 | 0.065868 | 0.07461 | 0.087194 | 0.10263 | 0.12083
| Normal | Default | 0.063429 | 0.068744 | 0.079952 | 0.091303 | 0.10753 | 0.12779 | 0.15343
Head Clarity | Low | Default | 0.12118 | 0.13092 | 0.14699 | 0.16349 | 0.18692 | 0.21669 | 0.25525
| Medium | Default | 0.18624 | 0.20175 | 0.23416 | 0.26727 | 0.3157 | 0.36742 | 0.42739
| Normal | Default | 0.34284 | 0.37 | 0.42562 | 0.48177 | 0.56329 | 0.66105 | 0.78392
Spine Clarity | Low | Default | 0.12118 | 0.13092 | 0.14699 | 0.16349 | 0.18692 | 0.21669 | 0.25525
| Medium | Default | 0.18624 | 0.20175 | 0.23416 | 0.26727 | 0.3157 | 0.36742 | 0.42739
| Normal | Default | 0.34284 | 0.37 | 0.42562 | 0.48177 | 0.56329 | 0.66105 | 0.78392
Thorax Clarity | Low | Default | 0.12118 | 0.13092 | 0.14699 | 0.16349 | 0.18692 | 0.21669 | 0.25525
| Medium | Default | 0.18624 | 0.20175 | 0.23416 | 0.26727 | 0.3157 | 0.36742 | 0.42739
| Normal | Default | 0.34284 | 0.37 | 0.42562 | 0.48177 | 0.56329 | 0.66105 | 0.78392
Abdomen Clarity | Low | Default | 0.12118 | 0.13092 | 0.14699 | 0.16349 | 0.18692 | 0.21669 | 0.25525
| Medium | Default | 0.18624 | 0.20175 | 0.23416 | 0.26727 | 0.3157 | 0.36742 | 0.42739
| Normal | Default | 0.34284 | 0.37 | 0.42562 | 0.48177 | 0.56329 | 0.66105 | 0.78392
Peripheral Clarity | Low | Default | 0.090185 | 0.097607 | 0.11246 | 0.12729 | 0.14873 | 0.17472 | 0.20576
| Medium | Default | 0.13178 | 0.14246 | 0.16456 | 0.1869 | 0.21962 | 0.25926 | 0.31027
| Normal | Default | 0.3217 | 0.3462 | 0.39679 | 0.44665 | 0.51734 | 0.60362 | 0.7102
Radiography | Exposure | typical and worst case samples per application group
|  | field of View | 0 | 1 | 2 | 3 | 4 | 5 | 6
|  | Field Size (cm): | 39 | 37 | 31 | 27 | 22 | 19 | 15
| X-ray Protocol | Patient Type | 0: Ref. AK (mGy/image)
Cardio Clarity | Left Coronary 15 fps Low | Default | 0.022659 | 0.024606 | 0.028579 | 0.032659 | 0.038482 | 0.045456 | 0.054428
| Rotational Scan Prop Ang0 4s | Default | 0.12471 | 0.13458 | 0.15502 | 0.17486 | 0.20424 | 0.23799 | 0.28017
|  | Infant | 0.038864 | 0.041554 | 0.047067 | 0.052862 | 0.061037 | 0.071417 | 0.085174
|  | Neonate | 0.023759 | 0.025702 | 0.029859 | 0.034238 | 0.040878 | 0.049592 | 0.061013
Cardio Pediatrics Clarity <40 kg | 15 fps Contrast low | Default | 0.01166 | 0.012685 | 0.014716 | 0.016722 | 0.019655 | 0.023285 | 0.027796
| 15 fps Contrast normal | Default | 0.018338 | 0.019885 | 0.02328 | 0.02676 | 0.031844 | 0.038133 | 0.046095
Cardio Pediatrics Clarity >40 kg | 15 fps Contrast low | Default | 0.022659 | 0.024606 | 0.028579 | 0.032659 | 0.038482 | 0.045456 | 0.054428
|  | Neonate,Infant,Child | 0.01166 | 0.012685 | 0.014716 | 0.016722 | 0.019655 | 0.023285 | 0.027796
| 15 fps Contrast normal | Default | 0.048226 | 0.052371 | 0.0604 | 0.068728 | 0.080553 | 0.09445 | 0.11208
|  | Neonate,Infant,Child | 0.018338 | 0.019885 | 0.02328 | 0.02676 | 0.031844 | 0.038133 | 0.046095
EP Clarity | 15 Fps | Default | 0.022659 | 0.024606 | 0.028579 | 0.032659 | 0.038482 | 0.045456 | 0.054428
EP Mapping Clarity | 7.5 fps | Default | 0.022685 | 0.024594 | 0.028596 | 0.032639 | 0.038463 | 0.045492 | 0.054476
Head Clarity | Cerebral 2fps low | Default | 0.57309 | 0.63117 | 0.75391 | 0.88483 | 1.0756 | 1.3148 | 1.6248
| Aortic Arch 3 fps | Default | 0.81733 | 0.8997 | 1.0757 | 1.2605 | 1.5342 | 1.8757 | 2.319
Spine Clarity | 4 fps | Default | 0.49225 | 0.54219 | 0.65071 | 0.76584 | 0.93323 | 1.1496 | 1.4285
| 2 fps | Default | 0.49198 | 0.54292 | 0.65047 | 0.76505 | 0.93369 | 1.1489 | 1.4283
Thorax Clarity | Lungs 2 fps | Default | 0.70017 | 0.77203 | 0.92352 | 1.0846 | 1.3214 | 1.6198 | 2.0099
| Subclavian 3 fps | Default | 0.98835 | 1.09 | 1.3006 | 1.5226 | 1.8476 | 2.2526 | 2.7727
Abdomen Clarity | 3 fps low | Default | 0.45973 | 0.50697 | 0.60543 | 0.70921 | 0.8614 | 1.0513 | 1.2987
| Iliac / Pelvis 3fps | Default | 2.2642 | 2.4997 | 2.9909 | 3.5199 | 4.2904 | 5.2618 | 6.5381
Peripheral Clarity | Upper Legs 3 fps | Default | 0.64145 | 0.70574 | 0.82749 | 0.9231 | 1.0593 | 1.2161 | 1.3896
| Lower Legs 1 fps | Default | 1.1836 | 1.3029 | 1.5531 | 1.812 | 2.1316 | 2.4021 | 2.7299
Roadmap |  | All modes for all application groups
|  | field of View | 0 | 1 | 2 | 3 | 4 | 5 | 6
|  | Field Size (cm): | 39 | 37 | 31 | 27 | 22 | 19 | 15
| mode | Patient Type | 0: Ref. AK (mGy/image)
Cardio Clarity | Navigate | Default | 0.096408 | 0.10314 | 0.11684 | 0.1304 | 0.15011 | 0.17353 | 0.20286
| UnSubtract | Default | 0.096408 | 0.10314 | 0.11684 | 0.1304 | 0.15011 | 0.17353 | 0.20286
Cardio Pediatrics Clarity <40 kg | Navigate | Default | 0.096408 | 0.10314 | 0.11684 | 0.1304 | 0.15011 | 0.17353 | 0.20286
| UnSubtract | Default | 0.096408 | 0.10314 | 0.11684 | 0.1304 | 0.15011 | 0.17353 | 0.20286
Cardio Pediatrics Clarity >40 kg | Navigate | Default | 0.096408 | 0.10314 | 0.11684 | 0.1304 | 0.15011 | 0.17353 | 0.20286
| UnSubtract | Default | 0.096408 | 0.10314 | 0.11684 | 0.1304 | 0.15011 | 0.17353 | 0.20286
EP Clarity | Navigate | Default | 0.096408 | 0.10314 | 0.11684 | 0.1304 | 0.15011 | 0.17353 | 0.20286
| UnSubtract | Default | 0.096408 | 0.10314 | 0.11684 | 0.1304 | 0.15011 | 0.17353 | 0.20286
EP Mapping Clarity | Navigate | Default | 0.096408 | 0.10314 | 0.11684 | 0.1304 | 0.15011 | 0.17353 | 0.20286
| UnSubtract | Default | 0.096408 | 0.10314 | 0.11684 | 0.1304 | 0.15011 | 0.17353 | 0.20286
Head Clarity | Navigate | Default | 0.096408 | 0.10314 | 0.11684 | 0.1304 | 0.15011 | 0.17353 | 0.20286
| Coil | Default | 0.19205 | 0.21144 | 0.25266 | 0.29641 | 0.36003 | 0.44091 | 0.5444
| UnSubtract | Default | 0.096408 | 0.10314 | 0.11684 | 0.1304 | 0.15011 | 0.17353 | 0.20286
Spine Clarity | Navigate | Default | 0.096408 | 0.10314 | 0.11684 | 0.1304 | 0.15011 | 0.17353 | 0.20286
| Coil | Default | 0.14535 | 0.16033 | 0.19184 | 0.22481 | 0.27294 | 0.33416 | 0.41254
| UnSubtract | Default | 0.096408 | 0.10314 | 0.11684 | 0.1304 | 0.15011 | 0.17353 | 0.20286
Thorax Clarity | UnSubtract | Default | 0.096408 | 0.10314 | 0.11684 | 0.1304 | 0.15011 | 0.17353 | 0.20286
| Navigate | Default | 0.096408 | 0.10314 | 0.11684 | 0.1304 | 0.15011 | 0.17353 | 0.20286
Abdomen Clarity | UnSubtract | Default | 0.096408 | 0.10314 | 0.11684 | 0.1304 | 0.15011 | 0.17353 | 0.20286
| Navigate | Default | 0.096408 | 0.10314 | 0.11684 | 0.1304 | 0.15011 | 0.17353 | 0.20286
Peripheral Clarity | Navigate | Default | 0.096408 | 0.10314 | 0.11684 | 0.1304 | 0.15011 | 0.17353 | 0.20286
| UnSubtract | Default | 0.096408 | 0.10314 | 0.11684 | 0.1304 | 0.15011 | 0.17353 | 0.20286

## FD20
ch 16.20.5 | C20/F20 Systems
ch 16.20.5.1 | RAK values | 200 mm PMMA | Large SID
| Radioscopy | (Fluoro) | All modes for all application groups
|  | Field of View | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7
|  | Field Size (cm) | 48 | 42 | 37 | 31 | 27 | 22 | 19 | 15
| Flavor | Patient Type | All modes for all application groups 0: Ref. AK (mGy/s)
Cardio | Low | Default | 0.1021 | 0.12933 | 0.14276 | 0.16571 | 0.19311 | 0.23007 | 0.26772 | 0.3277
| Norm | Default | 0.30245 | 0.36716 | 0.39854 | 0.45055 | 0.51043 | 0.58788 | 0.66433 | 0.78364
| High | Default | 0.3725 | 0.45815 | 0.5009 | 0.56914 | 0.64927 | 0.75399 | 0.85688 | 1.0174
Pediatrics | Low | Default | 0.096027 | 0.11921 | 0.13084 | 0.15039 | 0.1735 | 0.20478 | 0.23671 | 0.2727
| Norm | Default | 0.217 | 0.26904 | 0.2942 | 0.33563 | 0.38584 | 0.45042 | 0.50265 | 0.58355
| High | Default | 0.38444 | 0.46276 | 0.49922 | 0.55836 | 0.62484 | 0.68821 | 0.74328 | 0.82573
EP | Low | Default | 0.047059 | 0.059545 | 0.065712 | 0.074297 | 0.085141 | 0.098291 | 0.11462 | 0.12761
| Norm | Default | 0.10423 | 0.13169 | 0.1458 | 0.16671 | 0.19965 | 0.2257 | 0.27678 | 0.31745
| High | Default | 0.12342 | 0.15655 | 0.1728 | 0.20052 | 0.23401 | 0.27867 | 0.32385 | 0.39678
EP Mapping | Low | Default | 0.023521 | 0.029754 | 0.032882 | 0.037147 | 0.04253 | 0.049112 | 0.058029 | 0.065194
| Norm | Default | 0.052103 | 0.065912 | 0.072811 | 0.083432 | 0.099783 | 0.11303 | 0.13835 | 0.15884
| High | Default | 0.061724 | 0.078186 | 0.086431 | 0.10026 | 0.11698 | 0.13917 | 0.16197 | 0.19811
Head | Low | Default | 0.11281 | 0.14132 | 0.15527 | 0.17832 | 0.20634 | 0.23582 | 0.26663 | 0.29894
| Norm | Default | 0.24433 | 0.30124 | 0.32911 | 0.37579 | 0.44779 | 0.53245 | 0.5944 | 0.66171
| High | Default | 0.40555 | 0.49542 | 0.54081 | 0.61452 | 0.70063 | 0.81383 | 0.92722 | 1.1054
Spine | Low | Default | 0.11281 | 0.14132 | 0.15527 | 0.17832 | 0.20634 | 0.23582 | 0.26663 | 0.29894
| Norm | Default | 0.24433 | 0.30124 | 0.32911 | 0.37579 | 0.44779 | 0.53245 | 0.5944 | 0.66171
| High | Default | 0.40555 | 0.49542 | 0.54081 | 0.61452 | 0.70063 | 0.81383 | 0.92722 | 1.1054
Thorax | Low | Default | 0.11281 | 0.14132 | 0.15527 | 0.17832 | 0.20634 | 0.23582 | 0.26663 | 0.29894
| Norm | Default | 0.24433 | 0.30124 | 0.32911 | 0.37579 | 0.44779 | 0.53245 | 0.5944 | 0.66171
| High | Default | 0.40555 | 0.49542 | 0.54081 | 0.61452 | 0.70063 | 0.81383 | 0.92722 | 1.1054
Abdomen | Low | Default | 0.11281 | 0.14132 | 0.15527 | 0.17832 | 0.20634 | 0.23582 | 0.26663 | 0.29894
| Norm | Default | 0.24433 | 0.30124 | 0.32911 | 0.37579 | 0.44779 | 0.53245 | 0.5944 | 0.66171
| High | Default | 0.40555 | 0.49542 | 0.54081 | 0.61452 | 0.70063 | 0.81383 | 0.92722 | 1.1054
Peripheral | Low | Default | 0.11281 | 0.14132 | 0.15527 | 0.17832 | 0.20634 | 0.23582 | 0.26663 | 0.29894
| Norm | Default | 0.24433 | 0.30124 | 0.32911 | 0.37579 | 0.44779 | 0.53245 | 0.5944 | 0.66171
| High | Default | 0.40555 | 0.49542 | 0.54081 | 0.61452 | 0.70063 | 0.81383 | 0.92722 | 1.1054
Radiography | Exposure | typical and worst case samples per application group
|  | field of View | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7
|  | Field Size (cm): | 48 | 42 | 37 | 31 | 27 | 22 | 19 | 15
| X-ray Protocol | Patient Type | 0: Ref. AK (mGy/image)
Cardio | Left Coronary | Default | 0.05329 | 0.066479 | 0.072982 | 0.083402 | 0.096084 | 0.11192 | 0.12823 | 0.15283
| 15 fps Low | Infant | 0.039418 | 0.048398 | 0.051885 | 0.057263 | 0.063537 | 0.071765 | 0.080113 | 0.093236
|  | Large Adult, VeryLarge Adult | 0.052421 | 0.065468 | 0.071846 | 0.082133 | 0.094605 | 0.11059 | 0.12616 | 0.15095
|  | Neonate | 0.020557 | 0.025412 | 0.027837 | 0.031948 | 0.037061 | 0.044053 | 0.051535 | 0.063547
| Rotational Scan | Default | 0.10293 | 0.12826 | 0.14006 | 0.16026 | 0.18396 | 0.21431 | 0.24437 | 0.29061
| Prop Ang0 4s | Infant | 0.034203 | 0.040983 | 0.044297 | 0.049789 | 0.056317 | 0.06505 | 0.074167 | 0.088504
|  | Neonate | 0.020557 | 0.025412 | 0.027837 | 0.031948 | 0.037061 | 0.044053 | 0.051535 | 0.063547
Pediatrics | 15 fps Contrast Normal | Default | 0.10505 | 0.13029 | 0.14228 | 0.1627 | 0.18639 | 0.21684 | 0.24704 | 0.29285
|  | Child,SmallAdult | 0.060312 | 0.075337 | 0.082641 | 0.094562 | 0.10866 | 0.12681 | 0.145 | 0.17316
|  | Infant | 0.034203 | 0.040983 | 0.044297 | 0.049789 | 0.056317 | 0.06505 | 0.074167 | 0.088504
|  | Neonate | 0.020557 | 0.025412 | 0.027837 | 0.031948 | 0.037061 | 0.044053 | 0.051535 | 0.063547
| Prop Free Position | Default | 0.10137 | 0.12679 | 0.13888 | 0.15899 | 0.1824 | 0.21265 | 0.24307 | 0.28919
|  | Neonate,Infant | 0.023134 | 0.028118 | 0.030494 | 0.034448 | 0.039191 | 0.045526 | 0.052136 | 0.062896
EP | 15 Fps | Default | 0.060312 | 0.075337 | 0.082641 | 0.094562 | 0.10866 | 0.12681 | 0.145 | 0.17316
| Prop 4s | Default | 0.11515 | 0.14724 | 0.18428 | 0.18599 | 0.18771 | 0.19005 | 0.19276 | 0.19578
|  | Neonate,Infant | 0.073093 | 0.094718 | 0.1202 | 0.12138 | 0.12276 | 0.12461 | 0.12691 | 0.12938
|  | LargeAdult,VeryLargeAdult | 0.11515 | 0.14724 | 0.18428 | 0.18599 | 0.18771 | 0.19005 | 0.19276 | 0.19578
EP Mapping | 7.5 fps low | Default | 0.00905 | 0.011388 | 0.01253 | 0.014095 | 0.016056 | 0.018444 | 0.021656 | 0.024208
| 3D EP Prop 4s | Default | 0.11515 | 0.14724 | 0.18428 | 0.18599 | 0.18771 | 0.19005 | 0.19276 | 0.19578
|  | Neonate,Infant | 0.073093 | 0.094718 | 0.1202 | 0.12138 | 0.12276 | 0.12461 | 0.12691 | 0.12938
|  | LargeAdult,VeryLargeAdult | 0.11515 | 0.14724 | 0.18428 | 0.18599 | 0.18771 | 0.19005 | 0.19276 | 0.19578
Head | Cerebral 2fps normal | Default | 2.2953 | 3.0037 | 3.3594 | 3.9729 | 4.7185 | 5.7314 | 6.7791 | 8.1027
|  | Child | 1.5253 | 1.8995 | 2.0444 | 2.2825 | 2.5652 | 2.9445 | 3.3541 | 3.8839
|  | Neonate,Infant | 1.6026 | 1.8995 | 2.0444 | 2.2825 | 2.5652 | 2.9445 | 3.3541 | 3.8839
| Aortic Arch 3 fps | Default | 3.5049 | 4.5851 | 5.1265 | 6.0606 | 7.0474 | 7.1207 | 7.2096 | 7.2992
Spine | 4 fps | Default | 0.96323 | 1.2587 | 1.4081 | 1.6642 | 1.9765 | 2.4034 | 2.8421 | 3.554
| 2 fps | Default | 1.203 | 1.572 | 1.759 | 2.0819 | 2.4697 | 3.0032 | 3.5547 | 4.4448
Thorax | Lungs 2 fps | Default | 1.1483 | 1.5018 | 1.6817 | 1.987 | 2.3568 | 2.7268 | 2.9623 | 3.2451
| Subclavian 3 fps | Default | 4.3185 | 5.6454 | 6.3185 | 7.4556 | 8.8298 | 10.7056 | 11.7468 | 12.5537
Abdomen | Abdomen | Default | 2.1399 | 2.7993 | 3.1307 | 3.6992 | 4.3887 | 5.3352 | 6.3039 | 6.3946
| 3 fps | LargeAdult, VeryLargeAdult | 1.8565 | 2.4313 | 2.7186 | 3.2144 | 3.8168 | 4.643 | 5.4958 | 6.8768
| Iliac / Pelvis 3fps | Default | 3.4763 | 4.5459 | 5.0889 | 6.0099 | 7.1382 | 8.6691 | 10.2551 | 12.2547
Peripheral | Upper Legs 3 fps | Default | 3.0689 | 4.0044 | 4.4825 | 5.2973 | 5.8448 | 6.3704 | 6.8689 | 7.481
| Lower Legs 1 fps | Default | 3.8102 | 4.9809 | 5.5574 | 6.1088 | 6.5679 | 7.1437 | 7.6877 | 8.3354
Roadmap |  | All modes for all application groups
|  | field of View | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7
|  | Field Size (cm): | 48 | 42 | 37 | 31 | 27 | 22 | 19 | 15
| mode | Patient Type | 0: Ref. AK (mGy/image)
Cardio | Navigate | Default | 0.20704 | 0.25682 | 0.2803 | 0.3196 | 0.36586 | 0.42322 | 0.45586 | 0.45969
| UnSubtract | Default | 0.20704 | 0.25682 | 0.2803 | 0.3196 | 0.36586 | 0.42322 | 0.45586 | 0.45969
Pediatrics | Navigate | Default | 0.20704 | 0.25682 | 0.2803 | 0.3196 | 0.36586 | 0.42322 | 0.45586 | 0.45969
| UnSubtract | Default | 0.20704 | 0.25682 | 0.2803 | 0.3196 | 0.36586 | 0.42322 | 0.45586 | 0.45969
EP | Navigate | Default | 0.20704 | 0.25682 | 0.2803 | 0.3196 | 0.36586 | 0.42322 | 0.45586 | 0.45969
EP Mapping | Navigate | Default | 0.20704 | 0.25682 | 0.2803 | 0.3196 | 0.36586 | 0.42322 | 0.45586 | 0.45969
Head | Navigate | Default | 0.20704 | 0.25682 | 0.2803 | 0.3196 | 0.36586 | 0.42322 | 0.45586 | 0.45969
| Carotid | Default | 0.18391 | 0.21922 | 0.23595 | 0.26279 | 0.2936 | 0.33382 | 0.35487 | 0.35834
| Coil | Default | 0.86853 | 1.1368 | 1.273 | 1.5038 | 1.784 | 2.1667 | 2.5595 | 3.203
Spine | UnSubtract | Default | 0.20704 | 0.25682 | 0.2803 | 0.3196 | 0.36586 | 0.42322 | 0.45586 | 0.45969
| Navigate | Default | 0.20704 | 0.25682 | 0.2803 | 0.3196 | 0.36586 | 0.42322 | 0.45586 | 0.45969
| Stent | Default | 0.18391 | 0.21922 | 0.23595 | 0.26279 | 0.2936 | 0.33382 | 0.35487 | 0.35834
Thorax | UnSubtract | Default | 0.20704 | 0.25682 | 0.2803 | 0.3196 | 0.36586 | 0.42322 | 0.45586 | 0.45969
| Navigate | Default | 0.20704 | 0.25682 | 0.2803 | 0.3196 | 0.36586 | 0.42322 | 0.45586 | 0.45969
| Intervention | Default | 0.18391 | 0.21922 | 0.23595 | 0.26279 | 0.2936 | 0.33382 | 0.35487 | 0.35834
Abdomen | UnSubtract | Default | 0.20704 | 0.25682 | 0.2803 | 0.3196 | 0.36586 | 0.42322 | 0.45586 | 0.45969
| Navigate | Default | 0.20704 | 0.25682 | 0.2803 | 0.3196 | 0.36586 | 0.42322 | 0.45586 | 0.45969
| Stent | Default | 0.18391 | 0.21922 | 0.23595 | 0.26279 | 0.2936 | 0.33382 | 0.35487 | 0.35834
Peripheral | UnSubtract | Default | 0.20704 | 0.25682 | 0.2803 | 0.3196 | 0.36586 | 0.42322 | 0.45586 | 0.45969
| Navigate | Default | 0.20704 | 0.25682 | 0.2803 | 0.3196 | 0.36586 | 0.42322 | 0.45586 | 0.45969

## FD20 Clarity
ch 16.20.6 | C20/F20 Systems | Clarity
ch 16.20.6.1 | RAK values | 200 mm PMMA | Large SID
| Radioscopy | (Fluoro) | All modes for all application groups
|  | Field of View | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7
|  | Field Size (cm) | 48 | 42 | 37 | 31 | 27 | 22 | 19 | 15
| Flavor | Patient Type | All modes for all application groups 0: Ref. AK (mGy/s)
Cardio Clarity | Low | Default | 0.0772 | 0.095593 | 0.10466 | 0.1192 | 0.13677 | 0.15976 | 0.18213 | 0.21609
| Medium | Default | 0.1147 | 0.14224 | 0.156 | 0.17856 | 0.20515 | 0.24063 | 0.27639 | 0.33319
| Normal | Default | 0.27669 | 0.3398 | 0.37004 | 0.42017 | 0.47803 | 0.55507 | 0.63015 | 0.74826
Cardio Pediatrics Clarity <40 kg | Low | Default | 0.032459 | 0.040248 | 0.044045 | 0.050505 | 0.058056 | 0.068272 | 0.07864 | 0.095067
| Medium | Default | 0.048201 | 0.060083 | 0.066059 | 0.075824 | 0.087409 | 0.10287 | 0.11872 | 0.14294
| Normal | Default | 0.074186 | 0.092607 | 0.10144 | 0.11651 | 0.13417 | 0.15734 | 0.17954 | 0.21427
Cardio Pediatrics Clarity >40 kg | Low | Default | 0.0772 | 0.095593 | 0.10466 | 0.1192 | 0.13677 | 0.15976 | 0.18213 | 0.21609
| Medium | Default | 0.1147 | 0.14224 | 0.156 | 0.17856 | 0.20515 | 0.24063 | 0.27639 | 0.33319
| Normal | Default | 0.27669 | 0.3398 | 0.37004 | 0.42017 | 0.47803 | 0.55507 | 0.63015 | 0.74826
EP Clarity | Low | Default | 0.029789 | 0.029938 | 0.030037 | 0.034192 | 0.039234 | 0.045874 | 0.052296 | 0.061909
| Medium | Default | 0.059025 | 0.059292 | 0.059469 | 0.067846 | 0.077842 | 0.091007 | 0.10393 | 0.12269
| Normal | Default | 0.075541 | 0.075903 | 0.076121 | 0.087248 | 0.10113 | 0.1193 | 0.13724 | 0.16599
EP Mapping Clarity | Low | Default | 0.029789 | 0.029938 | 0.030037 | 0.034192 | 0.039234 | 0.045874 | 0.052296 | 0.061909
| Medium | Default | 0.059025 | 0.059292 | 0.059469 | 0.067846 | 0.077842 | 0.091007 | 0.10393 | 0.12269
| Normal | Default | 0.075541 | 0.075903 | 0.076121 | 0.087248 | 0.10113 | 0.1193 | 0.13724 | 0.16599
Head Clarity | Low | Default | 0.10348 | 0.12827 | 0.1387 | 0.15456 | 0.17884 | 0.20897 | 0.23439 | 0.26427
| Medium | Default | 0.1569 | 0.1959 | 0.21514 | 0.2475 | 0.29735 | 0.35493 | 0.39548 | 0.44088
| Normal | Default | 0.2877 | 0.35639 | 0.38947 | 0.44562 | 0.51055 | 0.59491 | 0.681 | 0.81442
Spine Clarity | Low | Default | 0.10348 | 0.12827 | 0.1387 | 0.15456 | 0.17884 | 0.20897 | 0.23439 | 0.26427
| Medium | Default | 0.1569 | 0.1959 | 0.21514 | 0.2475 | 0.29735 | 0.35493 | 0.39548 | 0.44088
| Normal | Default | 0.2877 | 0.35639 | 0.38947 | 0.44562 | 0.51055 | 0.59491 | 0.681 | 0.81442
Thorax Clarity | Low | Default | 0.10348 | 0.12827 | 0.1387 | 0.15456 | 0.17884 | 0.20897 | 0.23439 | 0.26427
| Medium | Default | 0.1569 | 0.1959 | 0.21514 | 0.2475 | 0.29735 | 0.35493 | 0.39548 | 0.44088
| Normal | Default | 0.2877 | 0.35639 | 0.38947 | 0.44562 | 0.51055 | 0.59491 | 0.681 | 0.81442
Abdomen Clarity | Low | Default | 0.10348 | 0.12827 | 0.1387 | 0.15456 | 0.17884 | 0.20897 | 0.23439 | 0.26427
| Medium | Default | 0.1569 | 0.1959 | 0.21514 | 0.2475 | 0.29735 | 0.35493 | 0.39548 | 0.44088
| Normal | Default | 0.2877 | 0.35639 | 0.38947 | 0.44562 | 0.51055 | 0.59491 | 0.681 | 0.81442
Peripheral Clarity | Low | Default | 0.0772 | 0.095593 | 0.10466 | 0.1192 | 0.13677 | 0.15976 | 0.18213 | 0.21609
| Medium | Default | 0.1147 | 0.14224 | 0.156 | 0.17856 | 0.20515 | 0.24063 | 0.27639 | 0.33319
| Normal | Default | 0.27669 | 0.3398 | 0.37004 | 0.42017 | 0.47803 | 0.55507 | 0.63015 | 0.74826
Radiography | Exposure | typical and worst case samples per application group
|  | field of View | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7
|  | Field Size (cm): | 48 | 42 | 37 | 31 | 27 | 22 | 19 | 15
| X-ray Protocol | Patient Type | 0: Ref. AK (mGy/image)
Cardio Clarity | Left Coronary 15 fps Low | Default | 0.018683 | 0.023449 | 0.02581 | 0.029668 | 0.034334 | 0.040371 | 0.046559 | 0.056086
| Rotational Scan Prop Ang0 4s | Default | 0.10505 | 0.13023 | 0.14242 | 0.1627 | 0.18658 | 0.21685 | 0.24681 | 0.29292
|  | Infant | 0.034203 | 0.040983 | 0.044297 | 0.049789 | 0.056317 | 0.06505 | 0.074167 | 0.088504
|  | Neonate | 0.020557 | 0.025412 | 0.027837 | 0.031948 | 0.037061 | 0.044053 | 0.051535 | 0.063547
Cardio Pediatrics Clarity <40 kg | 15 fps Contrast low | Default | 0.009665 | 0.012219 | 0.01347 | 0.015374 | 0.017633 | 0.020636 | 0.023711 | 0.028415
| 15 fps Contrast normal | Default | 0.015286 | 0.019114 | 0.021063 | 0.024398 | 0.028368 | 0.033692 | 0.039087 | 0.047501
Cardio Pediatrics Clarity >40 kg | 15 fps Contrast low | Default | 0.018683 | 0.023449 | 0.02581 | 0.029668 | 0.034334 | 0.040371 | 0.046559 | 0.056086
|  | Neonate,Infant,Child | 0.009665 | 0.012219 | 0.01347 | 0.015374 | 0.017633 | 0.020636 | 0.023711 | 0.028415
| 15 fps Contrast normal | Default | 0.04088 | 0.05104 | 0.056094 | 0.064378 | 0.073837 | 0.086553 | 0.099036 | 0.11843
|  | Neonate,Infant,Child | 0.015286 | 0.019114 | 0.021063 | 0.024398 | 0.028368 | 0.033692 | 0.039087 | 0.047501
EP Clarity | 15 Fps | Default | 0.018683 | 0.023449 | 0.02581 | 0.029668 | 0.034334 | 0.040371 | 0.046559 | 0.056086
| Prop 4s | Default | 0.11515 | 0.14724 | 0.18428 | 0.18599 | 0.18771 | 0.19005 | 0.19276 | 0.19578
|  | Neonate,Infant | 0.073093 | 0.094718 | 0.1202 | 0.12138 | 0.12276 | 0.12461 | 0.12691 | 0.12938
|  | LargeAdult,VeryLargeAdult | 0.11515 | 0.14724 | 0.18428 | 0.18599 | 0.18771 | 0.19005 | 0.19276 | 0.19578
EP Mapping Clarity | 7.5 fps | Default | 0.018673 | 0.023468 | 0.025793 | 0.029675 | 0.034305 | 0.040381 | 0.046518 | 0.056053
| 3D EP Prop 4s | Default | 0.11515 | 0.14724 | 0.18428 | 0.18599 | 0.18771 | 0.19005 | 0.19276 | 0.19578
|  | Neonate,Infant | 0.073093 | 0.094718 | 0.1202 | 0.12138 | 0.12276 | 0.12461 | 0.12691 | 0.12938
|  | LargeAdult,VeryLargeAdult | 0.11515 | 0.14724 | 0.18428 | 0.18599 | 0.18771 | 0.19005 | 0.19276 | 0.19578
Head Clarity | Cerebral 2fps low | Default | 0.46598 | 0.60921 | 0.68194 | 0.80541 | 0.95655 | 1.1591 | 1.3704 | 1.6386
| Aortic Arch 3 fps | Default | 0.66557 | 0.87027 | 0.97499 | 1.1509 | 1.3647 | 1.6556 | 1.9587 | 2.3386
Spine Clarity | 4 fps | Default | 0.53402 | 0.69757 | 0.78097 | 0.92316 | 1.0984 | 1.3355 | 1.5855 | 1.8987
| 2 fps | Default | 0.53412 | 0.69863 | 0.78106 | 0.92331 | 1.0982 | 1.3363 | 1.5829 | 1.8984
Thorax Clarity | Lungs 2 fps | Default | 0.57403 | 0.75091 | 0.84076 | 0.99251 | 1.179 | 1.4327 | 1.6948 | 2.03
| Subclavian 3 fps | Default | 0.80148 | 1.0469 | 1.1721 | 1.3842 | 1.6408 | 1.9878 | 2.3448 | 2.7928
Abdomen Clarity | 3 fps low | Default | 0.60171 | 0.78718 | 0.88025 | 1.0403 | 1.2364 | 1.5014 | 1.7784 | 2.1255
| Iliac / Pelvis 3fps | Default | 0.57388 | 0.75101 | 0.84001 | 0.99309 | 1.1795 | 1.432 | 1.6949 | 2.0295
Peripheral Clarity | Upper Legs 3 fps | Default | 0.5208 | 0.68054 | 0.76146 | 0.86628 | 0.9755 | 1.1213 | 1.2499 | 1.3969
| Lower Legs 1 fps | Default | 0.95348 | 1.2451 | 1.3932 | 1.6419 | 1.9443 | 2.2271 | 2.4582 | 2.7395
Roadmap |  | All modes for all application groups
|  | field of View | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7
|  | Field Size (cm): | 48 | 42 | 37 | 31 | 27 | 22 | 19 | 15
| mode | Patient Type | 0: Ref. AK (mGy/image)
Cardio Clarity | Navigate | Default | 0.083224 | 0.1006 | 0.1086 | 0.12248 | 0.13808 | 0.15883 | 0.17902 | 0.2115
| UnSubtract | Default | 0.083224 | 0.1006 | 0.1086 | 0.12248 | 0.13808 | 0.15883 | 0.17902 | 0.2115
Cardio Pediatrics Clarity <40 kg | Navigate | Default | 0.083224 | 0.1006 | 0.1086 | 0.12248 | 0.13808 | 0.15883 | 0.17902 | 0.2115
| UnSubtract | Default | 0.083224 | 0.1006 | 0.1086 | 0.12248 | 0.13808 | 0.15883 | 0.17902 | 0.2115
Cardio Pediatrics Clarity >40 kg | Navigate | Default | 0.083224 | 0.1006 | 0.1086 | 0.12248 | 0.13808 | 0.15883 | 0.17902 | 0.2115
| UnSubtract | Default | 0.083224 | 0.1006 | 0.1086 | 0.12248 | 0.13808 | 0.15883 | 0.17902 | 0.2115
EP Clarity | Navigate | Default | 0.083224 | 0.1006 | 0.1086 | 0.12248 | 0.13808 | 0.15883 | 0.17902 | 0.2115
| UnSubtract | Default | 0.083224 | 0.1006 | 0.1086 | 0.12248 | 0.13808 | 0.15883 | 0.17902 | 0.2115
EP Mapping Clarity | Navigate | Default | 0.083224 | 0.1006 | 0.1086 | 0.12248 | 0.13808 | 0.15883 | 0.17902 | 0.2115
| UnSubtract | Default | 0.083224 | 0.1006 | 0.1086 | 0.12248 | 0.13808 | 0.15883 | 0.17902 | 0.2115
Head Clarity | Navigate | Default | 0.083224 | 0.1006 | 0.1086 | 0.12248 | 0.13808 | 0.15883 | 0.17902 | 0.2115
| Carotid | Default | 0.15637 | 0.20466 | 0.22907 | 0.27067 | 0.32117 | 0.38913 | 0.45983 | 0.57485
| Coil | Default | 0.083224 | 0.1006 | 0.1086 | 0.12248 | 0.13808 | 0.15883 | 0.17902 | 0.2115
Spine Clarity | UnSubtract | Default | 0.083224 | 0.1006 | 0.1086 | 0.12248 | 0.13808 | 0.15883 | 0.17902 | 0.2115
| Navigate | Default | 0.15637 | 0.20466 | 0.22907 | 0.27067 | 0.32117 | 0.38913 | 0.45983 | 0.57485
| Stent | Default | 0.083224 | 0.1006 | 0.1086 | 0.12248 | 0.13808 | 0.15883 | 0.17902 | 0.2115
Thorax Clarity | UnSubtract | Default | 0.083224 | 0.1006 | 0.1086 | 0.12248 | 0.13808 | 0.15883 | 0.17902 | 0.2115
| Navigate | Default | 0.083224 | 0.1006 | 0.1086 | 0.12248 | 0.13808 | 0.15883 | 0.17902 | 0.2115
Abdomen Clarity | UnSubtract | Default | 0.083224 | 0.1006 | 0.1086 | 0.12248 | 0.13808 | 0.15883 | 0.17902 | 0.2115
| Navigate | Default | 0.083224 | 0.1006 | 0.1086 | 0.12248 | 0.13808 | 0.15883 | 0.17902 | 0.2115
Peripheral Clarity | UnSubtract | Default | 0.083224 | 0.1006 | 0.1086 | 0.12248 | 0.13808 | 0.15883 | 0.17902 | 0.2115
| Navigate | Default | 0.083224 | 0.1006 | 0.1086 | 0.12248 | 0.13808 | 0.15883 | 0.17902 | 0.2115

## FD1212
ch 16.20.7 | FD1212 Systems
ch 16.20.7.1 | RAK values | 200 mm PMMA | Large SID
| Radioscopy | (Fluoro) | All modes for all application groups
|  | Field of View | 0 | 1 | 2 | 3 | 4 | 0 | 1 | 2 | 3 | 4
|  | Field Size (cm) | 30 | 27 | 22 | 19 | 15 | 30 | 27 | 22 | 19 | 15
| Flavor | Patient Type | Ref. AK (mGy/s) Frontal |  |  |  |  | Ref. AK (mGy/s) Lateral
Cardio | Low | Default | 0.23301 | 0.25446 | 0.2963 | 0.34946 | 0.42012 | 0.26314 | 0.28658 | 0.33369 | 0.39497 | 0.47935
| Norm | Default | 0.51156 | 0.66403 | 0.84957 | 0.86087 | 0.87621 | 0.57586 | 0.74851 | 0.95675 | 0.96971 | 0.98668
| High | Default | 0.70902 | 0.75587 | 0.88446 | 1.0258 | 1.1829 | 0.80155 | 0.85048 | 0.99666 | 1.1552 | 1.3494
Pediatrics | Low | Default | 0.18264 | 0.19939 | 0.23308 | 0.27555 | 0.32833 | 0.20631 | 0.22531 | 0.26354 | 0.31123 | 0.37107
| Norm | Default | 0.43298 | 0.46974 | 0.53864 | 0.61075 | 0.69784 | 0.49956 | 0.54023 | 0.62955 | 0.72246 | 0.82473
| High | Default | 0.68428 | 0.73105 | 0.79223 | 0.86437 | 0.95107 | 0.77355 | 0.82961 | 0.9371 | 1.0432 | 1.1455
EP | Low | Default | 0.086842 | 0.095231 | 0.11216 | 0.13186 | 0.15659 | 0.098504 | 0.10833 | 0.12813 | 0.15334 | 0.18249
| Norm | Default | 0.20609 | 0.22448 | 0.26146 | 0.30813 | 0.36535 | 0.23182 | 0.25248 | 0.29433 | 0.34799 | 0.41542
| High | Default | 0.24697 | 0.26934 | 0.31304 | 0.36813 | 0.43639 | 0.27863 | 0.30342 | 0.35261 | 0.41558 | 0.493
EP Mapping | Low | Default | 0.043419 | 0.047711 | 0.056487 | 0.067544 | 0.081573 | 0.049227 | 0.054162 | 0.064076 | 0.076638 | 0.092535
| Norm | Default | 0.10293 | 0.11224 | 0.13069 | 0.15436 | 0.18405 | 0.11584 | 0.12619 | 0.14715 | 0.17383 | 0.20771
| High | Default | 0.12348 | 0.13478 | 0.15659 | 0.18394 | 0.21822 | 0.13928 | 0.15163 | 0.1763 | 0.20765 | 0.24655
Head | Low | Default | 0.18425 | 0.25211 | 0.29081 | 0.29939 | 0.31022 | 0.20892 | 0.28662 | 0.33219 | 0.34226 | 0.35546
| Norm | Default | 0.47158 | 0.50321 | 0.56472 | 0.64077 | 0.73399 | 0.54334 | 0.58137 | 0.65698 | 0.74981 | 0.85713
| High | Default | 0.75284 | 0.81494 | 0.93732 | 1.0941 | 1.2778 | 0.84807 | 0.91913 | 1.0583 | 1.2332 | 1.444
Spine | Low | Default | 0.18425 | 0.25211 | 0.29081 | 0.29939 | 0.31022 | 0.20892 | 0.28662 | 0.33219 | 0.34226 | 0.35546
| Norm | Default | 0.47158 | 0.50321 | 0.56472 | 0.64077 | 0.73399 | 0.54334 | 0.58137 | 0.65698 | 0.74981 | 0.85713
| High | Default | 0.75284 | 0.81494 | 0.93732 | 1.0941 | 1.2778 | 0.84807 | 0.91913 | 1.0583 | 1.2332 | 1.444
Thorax | Low | Default | 0.18425 | 0.25211 | 0.29081 | 0.29939 | 0.31022 | 0.20892 | 0.28662 | 0.33219 | 0.34226 | 0.35546
| Norm | Default | 0.47158 | 0.50321 | 0.56472 | 0.64077 | 0.73399 | 0.54334 | 0.58137 | 0.65698 | 0.74981 | 0.85713
| High | Default | 0.75284 | 0.81494 | 0.93732 | 1.0941 | 1.2778 | 0.84807 | 0.91913 | 1.0583 | 1.2332 | 1.444
Abdomen | Low | Default | 0.18425 | 0.25211 | 0.29081 | 0.29939 | 0.31022 | 0.20892 | 0.28662 | 0.33219 | 0.34226 | 0.35546
| Norm | Default | 0.47158 | 0.50321 | 0.56472 | 0.64077 | 0.73399 | 0.54334 | 0.58137 | 0.65698 | 0.74981 | 0.85713
| High | Default | 0.75284 | 0.81494 | 0.93732 | 1.0941 | 1.2778 | 0.84807 | 0.91913 | 1.0583 | 1.2332 | 1.444
Peripheral | Low | Default | 0.18425 | 0.25211 | 0.29081 | 0.29939 | 0.31022 | 0.20892 | 0.28662 | 0.33219 | 0.34226 | 0.35546
| Norm | Default | 0.47158 | 0.50321 | 0.56472 | 0.64077 | 0.73399 | 0.54334 | 0.58137 | 0.65698 | 0.74981 | 0.85713
| High | Default | 0.75284 | 0.81494 | 0.93732 | 1.0941 | 1.2778 | 0.84807 | 0.91913 | 1.0583 | 1.2332 | 1.444
Radiography | Exposure | typical and worst case samples per application group
|  | field of View | 0 | 1 | 2 | 3 | 4 | 0 | 1 | 2 | 3 | 4
|  | Field Size (cm): | 30 | 27 | 22 | 19 | 15 | 30 | 27 | 22 | 19 | 15
| X-ray Protocol | Patient Type | Ref. AK (mGy/image) Frontal |  |  |  |  | Ref. AK (mGy/image) Lateral
Cardio | Left Coronary | Default | 0.10377 | 0.11258 | 0.13029 | 0.15187 | 0.17787 | 0.11732 | 0.12766 | 0.14735 | 0.17172 | 0.20128
| 15 fps Low | Infant | 0.073805 | 0.078544 | 0.087842 | 0.099428 | 0.11353 | 0.081377 | 0.086682 | 0.097077 | 0.11001 | 0.12592
|  | Large Adult, VeryLarge Adult | 0.10369 | 0.11263 | 0.13037 | 0.15192 | 0.17797 | 0.1173 | 0.1277 | 0.14738 | 0.17172 | 0.20136
|  | Neonate | 0.040121 | 0.043819 | 0.051496 | 0.061546 | 0.074395 | 0.048368 | 0.052427 | 0.060644 | 0.071575 | 0.085721
| Rotational Scan | Default | 0.19519 | 0.21181 | 0.24373 | 0.28212 | 0.32934 | - | - | - | - | -
| Prop Ang0 4s | Infant | 0.063269 | 0.068207 | 0.077879 | 0.090131 | 0.10568 | - | - | - | - | -
|  | Neonate | 0.040121 | 0.043819 | 0.051496 | 0.061546 | 0.074395 | - | - | - | - | -
Pediatrics | 15 fps Contrast Normal | Default | 0.19516 | 0.21185 | 0.24377 | 0.28211 | 0.32928 | 0.22168 | 0.24028 | 0.27566 | 0.32028 | 0.37216
|  | Child,SmallAdult | 0.11548 | 0.12521 | 0.1449 | 0.16912 | 0.19809 | 0.13059 | 0.14194 | 0.16418 | 0.19137 | 0.22381
|  | Infant | 0.073805 | 0.078544 | 0.087842 | 0.099428 | 0.11353 | 0.081377 | 0.086682 | 0.097077 | 0.11001 | 0.12592
|  | Neonate | 0.043334 | 0.046926 | 0.054155 | 0.063591 | 0.075924 | 0.048368 | 0.052427 | 0.060644 | 0.071575 | 0.085721
| Prop Free Position | Default | 0.20113 | 0.21835 | 0.25218 | 0.29298 | 0.34168 | - | - | - | - | -
|  | Neonate,Infant | 0.043334 | 0.046926 | 0.054155 | 0.063591 | 0.075924 | - | - | - | - | -
EP | 15 Fps | Default | 0.11548 | 0.12521 | 0.1449 | 0.16912 | 0.19809 | 0.13059 | 0.14194 | 0.16418 | 0.19137 | 0.22381
| Prop 4s | Default | 0.22823 | 0.23056 | 0.2346 | 0.23976 | 0.24627 | - | - | - | - | -
|  | Neonate,Infant | 0.14768 | 0.14955 | 0.15242 | 0.1563 | 0.16075 | - | - | - | - | -
|  | LargeAdult,VeryLargeAdult | 0.22823 | 0.23056 | 0.2346 | 0.23976 | 0.24627 | - | - | - | - | -
EP Mapping | 7.5 fps low | Default | 0.016591 | 0.018228 | 0.021467 | 0.025608 | 0.030723 | 0.018862 | 0.020717 | 0.024386 | 0.029091 | 0.034926
| 3D EP Prop 4s | Default | 0.22823 | 0.23056 | 0.2346 | 0.23976 | 0.24627 | - | - | - | - | -
|  | Neonate,Infant | 0.14768 | 0.14955 | 0.15242 | 0.1563 | 0.16075 | - | - | - | - | -
|  | LargeAdult,VeryLargeAdult | 0.22823 | 0.23056 | 0.2346 | 0.23976 | 0.24627 | - | - | - | - | -
Head | Cerebral 2fps normal | Default | 4.895 | 5.436 | 6.5402 | 7.7015 | 8.5024 | 5.6535 | 6.2785 | 7.5538 | 8.3908 | 9.2806
|  | Child | 3.1525 | 3.3776 | 3.8201 | 4.3787 | 5.0724 | 3.4913 | 3.7392 | 4.2334 | 4.8597 | 5.6551
|  | Neonate,Infant | 3.1525 | 3.3776 | 3.8201 | 4.3787 | 5.0724 | 3.4913 | 3.7392 | 4.2334 | 4.8597 | 5.6551
| Aortic Arch 3 fps | Default | 7.3036 | 8.1135 | 9.7134 | 11.7876 | 12.8945 | 8.4783 | 9.3995 | 11.2761 | 13.0289 | 13.8561
Spine | 4 fps | Default | 2.0496 | 2.2788 | 2.7391 | 3.3353 | 4.0877 | 1.8065 | 2.0114 | 2.4278 | 2.9667 | 3.648
| 2 fps | Default | 2.5655 | 2.8512 | 3.4278 | 4.1721 | 5.1126 | 2.2577 | 2.5169 | 3.0348 | 3.7113 | 4.5569
Thorax | Lungs 2 fps | Default | 2.446 | 2.7185 | 3.2691 | 3.9505 | 4.361 | 2.8274 | 3.1425 | 3.779 | 4.3046 | 4.7604
| Subclavian 3 fps | Default | 9.046 | 10.6732 | 10.8101 | 10.9715 | 11.1828 | 10.5111 | 12.3857 | 12.5437 | 12.7317 | 12.977
Abdomen | Abdomen | Default | 3.8242 | 4.2416 | 5.1008 | 6.1983 | 7.5784 | 4.4295 | 4.9232 | 5.9083 | 7.1793 | 8.7779
| 3 fps | LargeAdult, VeryLargeAdult | 3.3166 | 3.681 | 4.4337 | 5.3978 | 6.6204 | 3.8422 | 4.2642 | 5.1308 | 6.253 | 7.653
| Iliac / Pelvis 3fps | Default | 7.2974 | 8.1082 | 9.7297 | 11.8418 | 14.4718 | 8.4486 | 9.3873 | 11.2874 | 13.71 | 16.7548
Peripheral | Upper Legs 3 fps | Default | 6.6081 | 7.3341 | 8.209 | 8.9328 | 9.7813 | 6.0482 | 6.3365 | 6.8937 | 7.5845 | 8.4128
| Lower Legs 1 fps | Default | 8.2287 | 8.6443 | 9.3197 | 10.1247 | 11.0434 | 6.7901 | 7.1074 | 7.7072 | 8.4366 | 9.3184
Roadmap |  | All modes for all application groups
|  | field of View | 0 | 1 | 2 | 3 | 4 | 0 | 1 | 2 | 3 | 4
|  | Field Size (cm): | 30 | 27 | 22 | 19 | 15 | 30 | 27 | 22 | 19 | 15
| mode | Patient Type | Ref. AK (mGy/image) Frontal |  |  |  |  | Ref. AK (mGy/image) Lateral
Cardio | Navigate | Default | 0.38281 | 0.48589 | 0.49039 | 0.49568 | 0.50257 | 0.43361 | 0.54941 | 0.55447 | 0.56043 | 0.56906
| UnSubtract | Default | 0.38281 | 0.48589 | 0.49039 | 0.49568 | 0.50257 | 0.43361 | 0.54941 | 0.55447 | 0.56043 | 0.56906
Pediatrics | Navigate | Default | 0.38281 | 0.48589 | 0.49039 | 0.49568 | 0.50257 | 0.43361 | 0.54941 | 0.55447 | 0.56043 | 0.56906
| UnSubtract | Default | 0.38281 | 0.48589 | 0.49039 | 0.49568 | 0.50257 | 0.43361 | 0.54941 | 0.55447 | 0.56043 | 0.56906
EP | Navigate | Default | 0.38281 | 0.48589 | 0.49039 | 0.49568 | 0.50257 | 0.43361 | 0.54941 | 0.55447 | 0.56043 | 0.56906
EP Mapping | Navigate | Default | 0.38281 | 0.48589 | 0.49039 | 0.49568 | 0.50257 | 0.43361 | 0.54941 | 0.55447 | 0.56043 | 0.56906
Head | Navigate | Default | 0.38281 | 0.48589 | 0.49039 | 0.49568 | 0.50257 | 0.43361 | 0.54941 | 0.55447 | 0.56043 | 0.56906
| Carotid | Default | 0.38058 | 0.46141 | 0.46528 | 0.47036 | 0.47561 | 0.42327 | 0.51407 | 0.51743 | 0.52352 | 0.53081
| Coil | Default | 1.8177 | 2.0195 | 2.4279 | 2.9511 | 3.6108 | 2.1096 | 2.3438 | 2.8116 | 3.4175 | 4.1814
Spine | UnSubtract | Default | 0.38281 | 0.48589 | 0.49039 | 0.49568 | 0.50257 | 0.43361 | 0.54941 | 0.55447 | 0.56043 | 0.56906
| Navigate | Default | 0.38281 | 0.48589 | 0.49039 | 0.49568 | 0.50257 | 0.43361 | 0.54941 | 0.55447 | 0.56043 | 0.56906
| Stent | Default | 0.38058 | 0.46141 | 0.46528 | 0.47036 | 0.47561 | 0.42327 | 0.51407 | 0.51743 | 0.52352 | 0.53081
Thorax | UnSubtract | Default | 0.38281 | 0.48589 | 0.49039 | 0.49568 | 0.50257 | 0.43361 | 0.54941 | 0.55447 | 0.56043 | 0.56906
| Navigate | Default | 0.38281 | 0.48589 | 0.49039 | 0.49568 | 0.50257 | 0.43361 | 0.54941 | 0.55447 | 0.56043 | 0.56906
| Intervention | Default | 0.38058 | 0.46141 | 0.46528 | 0.47036 | 0.47561 | 0.42327 | 0.51407 | 0.51743 | 0.52352 | 0.53081
Abdomen | UnSubtract | Default | 0.38281 | 0.48589 | 0.49039 | 0.49568 | 0.50257 | 0.43361 | 0.54941 | 0.55447 | 0.56043 | 0.56906
| Navigate | Default | 0.38281 | 0.48589 | 0.49039 | 0.49568 | 0.50257 | 0.43361 | 0.54941 | 0.55447 | 0.56043 | 0.56906
| Stent | Default | 0.38058 | 0.46141 | 0.46528 | 0.47036 | 0.47561 | 0.42327 | 0.51407 | 0.51743 | 0.52352 | 0.53081
Peripheral | UnSubtract | Default | 0.38281 | 0.48589 | 0.49039 | 0.49568 | 0.50257 | 0.43361 | 0.54941 | 0.55447 | 0.56043 | 0.56906
| Navigate | Default | 0.38281 | 0.48589 | 0.49039 | 0.49568 | 0.50257 | 0.43361 | 0.54941 | 0.55447 | 0.56043 | 0.56906

## FD1212 Clarity
ch 16.20.8 | FD1212 Systems | Clarity
ch 16.20.8.1 | RAK values | 200 mm PMMA | Large SID
| Radioscopy | (Fluoro) | All modes for all application groups
|  | Field of View | 0 | 1 | 2 | 3 | 4 | 0 | 1 | 2 | 3 | 4
|  | Field Size (cm) | 30 | 27 | 22 | 19 | 15 | 30 | 27 | 22 | 19 | 15
| Flavor | Patient Type | Ref. AK (mGy/s) Frontal |  |  |  |  | Ref. AK (mGy/s) Lateral
Cardio Clarity | Low | Default | 0.14706 | 0.15932 | 0.184 | 0.21366 | 0.24972 | 0.16536 | 0.17985 | 0.20754 | 0.24306 | 0.28578
| Medium | Default | 0.21664 | 0.23598 | 0.27409 | 0.32157 | 0.38007 | 0.24476 | 0.26662 | 0.30897 | 0.36327 | 0.42937
| Normal | Default | 0.51179 | 0.55279 | 0.63488 | 0.7371 | 0.85846 | 0.57549 | 0.62256 | 0.71522 | 0.8286 | 0.96791
Cardio Pediatrics Clarity <40 kg | Low | Default | 0.06093 | 0.066483 | 0.077564 | 0.091439 | 0.10864 | 0.068841 | 0.075114 | 0.087495 | 0.10315 | 0.12285
| Medium | Default | 0.092038 | 0.10026 | 0.11591 | 0.13527 | 0.15771 | 0.10406 | 0.11356 | 0.13212 | 0.15523 | 0.18077
| Normal | Default | 0.14401 | 0.15677 | 0.18098 | 0.21103 | 0.24769 | 0.16267 | 0.1768 | 0.20516 | 0.24057 | 0.2846
Cardio Pediatrics Clarity >40 kg | Low | Default | 0.14706 | 0.15932 | 0.184 | 0.21366 | 0.24972 | 0.16536 | 0.17985 | 0.20754 | 0.24306 | 0.28578
| Medium | Default | 0.21664 | 0.23598 | 0.27409 | 0.32157 | 0.38007 | 0.24476 | 0.26662 | 0.30897 | 0.36327 | 0.42937
| Normal | Default | 0.51179 | 0.55279 | 0.63488 | 0.7371 | 0.85846 | 0.57549 | 0.62256 | 0.71522 | 0.8286 | 0.96791
EP Clarity | Low | Default | 0.043606 | 0.047359 | 0.054727 | 0.063915 | 0.072768 | 0.049907 | 0.054199 | 0.062607 | 0.07312 | 0.086
| Medium | Default | 0.086443 | 0.093945 | 0.10857 | 0.12678 | 0.14713 | 0.10136 | 0.11006 | 0.12719 | 0.14868 | 0.17378
| Normal | Default | 0.10613 | 0.11606 | 0.13503 | 0.15906 | 0.18866 | 0.12026 | 0.13121 | 0.15289 | 0.18031 | 0.21362
EP Mapping Clarity | Low | Default | 0.043606 | 0.047359 | 0.054727 | 0.063915 | 0.072768 | 0.049907 | 0.054199 | 0.062607 | 0.07312 | 0.086
| Medium | Default | 0.086443 | 0.093945 | 0.10857 | 0.12678 | 0.14713 | 0.10136 | 0.11006 | 0.12719 | 0.14868 | 0.17378
| Normal | Default | 0.10613 | 0.11606 | 0.13503 | 0.15906 | 0.18866 | 0.12026 | 0.13121 | 0.15289 | 0.18031 | 0.21362
Head Clarity | Low | Default | 0.1886 | 0.20226 | 0.22932 | 0.26343 | 0.30709 | 0.21319 | 0.23027 | 0.26436 | 0.30769 | 0.36007
| Medium | Default | 0.30582 | 0.3337 | 0.38933 | 0.44716 | 0.51394 | 0.34573 | 0.37702 | 0.43598 | 0.50281 | 0.58525
| Normal | Default | 0.55654 | 0.60362 | 0.69734 | 0.81264 | 0.95099 | 0.62811 | 0.68181 | 0.78701 | 0.91775 | 1.0755
Spine Clarity | Low | Default | 0.1886 | 0.20226 | 0.22932 | 0.26343 | 0.30709 | 0.21319 | 0.23027 | 0.26436 | 0.30769 | 0.36007
| Medium | Default | 0.30582 | 0.3337 | 0.38933 | 0.44716 | 0.51394 | 0.34573 | 0.37702 | 0.43598 | 0.50281 | 0.58525
| Normal | Default | 0.55654 | 0.60362 | 0.69734 | 0.81264 | 0.95099 | 0.62811 | 0.68181 | 0.78701 | 0.91775 | 1.0755
Thorax Clarity | Low | Default | 0.1886 | 0.20226 | 0.22932 | 0.26343 | 0.30709 | 0.21319 | 0.23027 | 0.26436 | 0.30769 | 0.36007
| Medium | Default | 0.30582 | 0.3337 | 0.38933 | 0.44716 | 0.51394 | 0.34573 | 0.37702 | 0.43598 | 0.50281 | 0.58525
| Normal | Default | 0.55654 | 0.60362 | 0.69734 | 0.81264 | 0.95099 | 0.62811 | 0.68181 | 0.78701 | 0.91775 | 1.0755
Abdomen Clarity | Low | Default | 0.1886 | 0.20226 | 0.22932 | 0.26343 | 0.30709 | 0.21319 | 0.23027 | 0.26436 | 0.30769 | 0.36007
| Medium | Default | 0.30582 | 0.3337 | 0.38933 | 0.44716 | 0.51394 | 0.34573 | 0.37702 | 0.43598 | 0.50281 | 0.58525
| Normal | Default | 0.55654 | 0.60362 | 0.69734 | 0.81264 | 0.95099 | 0.62811 | 0.68181 | 0.78701 | 0.91775 | 1.0755
Peripheral Clarity | Low | Default | 0.1886 | 0.20226 | 0.22932 | 0.26343 | 0.30709 | 0.21319 | 0.23027 | 0.26436 | 0.30769 | 0.36007
| Medium | Default | 0.30582 | 0.3337 | 0.38933 | 0.44716 | 0.51394 | 0.34573 | 0.37702 | 0.43598 | 0.50281 | 0.58525
| Normal | Default | 0.55654 | 0.60362 | 0.69734 | 0.81264 | 0.95099 | 0.62811 | 0.68181 | 0.78701 | 0.91775 | 1.0755
Radiography | Exposure | typical and worst case samples per application group
|  | field of View | 0 | 1 | 2 | 3 | 4 | 0 | 1 | 2 | 3 | 4
|  | Field Size (cm): | 30 | 27 | 22 | 19 | 15 | 30 | 27 | 22 | 19 | 15
| X-ray Protocol | Patient Type | Ref. AK (mGy/image) Frontal |  |  |  |  | Ref. AK (mGy/image) Lateral
Cardio Clarity | Left Coronary 15 fps Low | Default | 0.036635 | 0.040075 | 0.046662 | 0.054835 | 0.064879 | 0.041601 | 0.045334 | 0.052813 | 0.062086 | 0.073439
| Rotational Scan Prop Ang0 4s | Default | 0.19519 | 0.21181 | 0.24373 | 0.28212 | 0.32934 | - | - | - | - | -
|  | Infant | 0.063269 | 0.068207 | 0.077879 | 0.090131 | 0.10568 | - | - | - | - | -
|  | Neonate | 0.040121 | 0.043819 | 0.051496 | 0.061546 | 0.074395 | - | - | - | - | -
Cardio Pediatrics Clarity <40 kg | 15 fps Contrast low | Default | 0.01854 | 0.020178 | 0.023419 | 0.027562 | 0.032418 | 0.02087 | 0.022698 | 0.026361 | 0.030975 | 0.036308
| 15 fps Contrast normal | Default | 0.030122 | 0.033021 | 0.038694 | 0.045989 | 0.054752 | 0.034243 | 0.037484 | 0.043974 | 0.052171 | 0.061111
Cardio Pediatrics Clarity >40 kg | 15 fps Contrast low | Default | 0.036635 | 0.040075 | 0.046662 | 0.054835 | 0.064879 | 0.041601 | 0.045334 | 0.052813 | 0.062086 | 0.073439
|  | Neonate,Infant,Child | 0.01854 | 0.020178 | 0.023419 | 0.027562 | 0.032418 | 0.02087 | 0.022698 | 0.026361 | 0.030975 | 0.036308
| 15 fps Contrast normal | Default | 0.077016 | 0.083838 | 0.096932 | 0.11321 | 0.13281 | 0.087402 | 0.095137 | 0.10995 | 0.12823 | 0.15008
|  | Neonate,Infant,Child | 0.030122 | 0.033021 | 0.038694 | 0.045989 | 0.054752 | 0.034243 | 0.037484 | 0.043974 | 0.052171 | 0.061111
EP Clarity | 15 Fps | Default | 0.036635 | 0.040075 | 0.046662 | 0.054835 | 0.064879 | 0.041601 | 0.045334 | 0.052813 | 0.062086 | 0.073439
| Prop 4s | Default | 0.22823 | 0.23056 | 0.2346 | 0.23976 | 0.24627 | - | - | - | - | -
|  | Neonate,Infant | 0.14768 | 0.14955 | 0.15242 | 0.1563 | 0.16075 | - | - | - | - | -
|  | LargeAdult,VeryLargeAdult | 0.22823 | 0.23056 | 0.2346 | 0.23976 | 0.24627 | - | - | - | - | -
EP Mapping Clarity | 7.5 fps | Default | 0.036629 | 0.040051 | 0.046635 | 0.054868 | 0.064863 | 0.041578 | 0.045289 | 0.052819 | 0.062101 | 0.073471
| 3D EP Prop 4s | Default | 0.22823 | 0.23056 | 0.2346 | 0.23976 | 0.24627 | - | - | - | - | -
|  | Neonate,Infant | 0.14768 | 0.14955 | 0.15242 | 0.1563 | 0.16075 | - | - | - | - | -
|  | LargeAdult,VeryLargeAdult | 0.22823 | 0.23056 | 0.2346 | 0.23976 | 0.24627 | - | - | - | - | -
Head Clarity | Cerebral 2fps low | Default | 0.99034 | 1.0986 | 1.3183 | 1.6019 | 1.9533 | 1.1432 | 1.2687 | 1.5237 | 1.8493 | 2.2584
| Aortic Arch 3 fps | Default | 1.4227 | 1.5797 | 1.8953 | 2.3051 | 2.8099 | 1.6455 | 1.8272 | 2.1922 | 2.6602 | 3.2502
Spine Clarity | 4 fps | Default | 1.1326 | 1.26 | 1.5194 | 1.8523 | 2.2728 | 1.2496 | 1.3872 | 1.6728 | 2.0392 | 2.5023
| 2 fps | Default | 1.1335 | 1.2601 | 1.5182 | 1.8526 | 2.2725 | 1.2484 | 1.3879 | 1.6721 | 2.0404 | 2.5004
Thorax Clarity | Lungs 2 fps | Default | 1.2209 | 1.3574 | 1.6298 | 1.9856 | 2.4316 | 1.4095 | 1.5671 | 1.8816 | 2.2923 | 2.8043
| Subclavian 3 fps | Default | 1.7217 | 1.9115 | 2.2886 | 2.7736 | 3.3778 | 1.9933 | 2.2105 | 2.6465 | 3.2033 | 3.901
Abdomen Clarity | 3 fps low | Default | 0.79655 | 0.88435 | 1.0596 | 1.2853 | 1.5686 | 0.92079 | 1.0222 | 1.2248 | 1.4857 | 1.8133
| Iliac / Pelvis 3fps | Default | 3.8904 | 4.3221 | 5.2019 | 6.3402 | 7.1376 | 3.9838 | 4.4346 | 5.3387 | 6.5205 | 7.354
Peripheral Clarity | Upper Legs 3 fps | Default | 1.128 | 1.2535 | 1.5015 | 1.6883 | 1.9032 | 1.3036 | 1.4453 | 1.6615 | 1.8615 | 2.1022
| Lower Legs 1 fps | Default | 2.0635 | 2.2854 | 2.7318 | 3.3018 | 4.0056 | 2.3908 | 2.6446 | 3.1613 | 3.8162 | 4.5026
Roadmap |  | All modes for all application groups
|  | field of View | 0 | 1 | 2 | 3 | 4 | 0 | 1 | 2 | 3 | 4
|  | Field Size (cm): | 30 | 27 | 22 | 19 | 15 | 30 | 27 | 22 | 19 | 15
| mode | Patient Type | Ref. AK (mGy/image) Frontal |  |  |  |  | Ref. AK (mGy/image) Lateral
Cardio Clarity | Navigate | Default | 0.17516 | 0.18801 | 0.21264 | 0.24323 | 0.27906 | 0.1956 | 0.20985 | 0.23782 | 0.27163 | 0.31263
| UnSubtract | Default | 0.17516 | 0.18801 | 0.21264 | 0.24323 | 0.27906 | 0.1956 | 0.20985 | 0.23782 | 0.27163 | 0.31263
Cardio Pediatrics Clarity <40 kg | Navigate | Default | 0.17516 | 0.18801 | 0.21264 | 0.24323 | 0.27906 | 0.1956 | 0.20985 | 0.23782 | 0.27163 | 0.31263
| UnSubtract | Default | 0.17516 | 0.18801 | 0.21264 | 0.24323 | 0.27906 | 0.1956 | 0.20985 | 0.23782 | 0.27163 | 0.31263
Cardio Pediatrics Clarity >40 kg | Navigate | Default | 0.17516 | 0.18801 | 0.21264 | 0.24323 | 0.27906 | 0.1956 | 0.20985 | 0.23782 | 0.27163 | 0.31263
| UnSubtract | Default | 0.17516 | 0.18801 | 0.21264 | 0.24323 | 0.27906 | 0.1956 | 0.20985 | 0.23782 | 0.27163 | 0.31263
EP Clarity | Navigate | Default | 0.17516 | 0.18801 | 0.21264 | 0.24323 | 0.27906 | 0.1956 | 0.20985 | 0.23782 | 0.27163 | 0.31263
| UnSubtract | Default | 0.17516 | 0.18801 | 0.21264 | 0.24323 | 0.27906 | 0.1956 | 0.20985 | 0.23782 | 0.27163 | 0.31263
EP Mapping Clarity | Navigate | Default | 0.17516 | 0.18801 | 0.21264 | 0.24323 | 0.27906 | 0.1956 | 0.20985 | 0.23782 | 0.27163 | 0.31263
| UnSubtract | Default | 0.17516 | 0.18801 | 0.21264 | 0.24323 | 0.27906 | 0.1956 | 0.20985 | 0.23782 | 0.27163 | 0.31263
Head Clarity | Navigate | Default | 0.17516 | 0.18801 | 0.21264 | 0.24323 | 0.27906 | 0.1956 | 0.20985 | 0.23782 | 0.27163 | 0.31263
| Carotid | Default | 0.33483 | 0.37173 | 0.44659 | 0.54229 | 0.6613 | 0.38748 | 0.42945 | 0.51588 | 0.62602 | 0.76429
| Coil | Default | 0.17516 | 0.18801 | 0.21264 | 0.24323 | 0.27906 | 0.1956 | 0.20985 | 0.23782 | 0.27163 | 0.31263
Spine Clarity | UnSubtract | Default | 0.17516 | 0.18801 | 0.21264 | 0.24323 | 0.27906 | 0.1956 | 0.20985 | 0.23782 | 0.27163 | 0.31263
| Navigate | Default | 0.2361 | 0.26218 | 0.31428 | 0.38168 | 0.46659 | 0.38748 | 0.42945 | 0.51588 | 0.62602 | 0.76429
| Stent | Default | 0.17516 | 0.18801 | 0.21264 | 0.24323 | 0.27906 | 0.1956 | 0.20985 | 0.23782 | 0.27163 | 0.31263
Thorax Clarity | UnSubtract | Default | 0.17516 | 0.18801 | 0.21264 | 0.24323 | 0.27906 | 0.1956 | 0.20985 | 0.23782 | 0.27163 | 0.31263
| Navigate | Default | 0.17516 | 0.18801 | 0.21264 | 0.24323 | 0.27906 | 0.1956 | 0.20985 | 0.23782 | 0.27163 | 0.31263
Abdomen Clarity | UnSubtract | Default | 0.17516 | 0.18801 | 0.21264 | 0.24323 | 0.27906 | 0.1956 | 0.20985 | 0.23782 | 0.27163 | 0.31263
| Navigate | Default | 0.17516 | 0.18801 | 0.21264 | 0.24323 | 0.27906 | 0.1956 | 0.20985 | 0.23782 | 0.27163 | 0.31263
Peripheral Clarity | UnSubtract | Default | 0.17516 | 0.18801 | 0.21264 | 0.24323 | 0.27906 | 0.1956 | 0.20985 | 0.23782 | 0.27163 | 0.31263
| Navigate | Default | 0.17516 | 0.18801 | 0.21264 | 0.24323 | 0.27906 | 0.1956 | 0.20985 | 0.23782 | 0.27163 | 0.31263

## FD2012
ch 16.20.9 | FD2012 Systems
ch 16.20.9.1 | RAK values | 200 mm PMMA | Large SID
| Radioscopy | (Fluoro) | All modes for all application groups
|  | Field of View | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 0 | 1 | 2 | 3 | 4
|  | Field Size (cm) | 48 | 42 | 37 | 31 | 27 | 22 | 19 | 15 | 30 | 27 | 22 | 19 | 15
| Flavor | Patient Type | All modes for all application groups 0: Ref. AK (mGy/s) Frontal | Ref. AK (mGy/s) Lateral
Cardio | Low | Default | 0.12462 | 0.15767 | 0.17404 | 0.20059 | 0.23236 | 0.27358 | 0.31434 | 0.37916 | 0.33692 | 0.36565 | 0.42766 | 0.5037 | 0.59367
| Norm | Default | 0.33822 | 0.41194 | 0.44658 | 0.5032 | 0.56729 | 0.6515 | 0.73118 | 0.84673 | 0.67235 | 0.87784 | 1.1181 | 1.1181 | 1.1181
| High | Default | 0.42296 | 0.52058 | 0.56779 | 0.64292 | 0.73021 | 0.84219 | 0.95071 | 1.1117 | 0.93641 | 0.98316 | 1.1516 | 1.3232 | 1.5245
Pediatrics | Low | Default | 0.11534 | 0.14388 | 0.15792 | 0.18096 | 0.20799 | 0.24305 | 0.2775 | 0.31485 | 0.26388 | 0.28661 | 0.33107 | 0.38516 | 0.44932
| Norm | Default | 0.25304 | 0.31318 | 0.34261 | 0.39007 | 0.44262 | 0.50052 | 0.55692 | 0.64479 | 0.60781 | 0.6424 | 0.73152 | 0.82659 | 0.93683
| High | Default | 0.42759 | 0.51388 | 0.55313 | 0.61486 | 0.66286 | 0.7243 | 0.78171 | 0.87065 | 0.8824 | 0.93314 | 1.0133 | 1.1077 | 1.2134
EP | Low | Default | 0.059046 | 0.074694 | 0.082365 | 0.092513 | 0.10445 | 0.118 | 0.13593 | 0.14932 | 0.13231 | 0.14442 | 0.16614 | 0.19277 | 0.22422
| Norm | Default | 0.12991 | 0.1623 | 0.17767 | 0.20307 | 0.24295 | 0.27265 | 0.33125 | 0.37483 | 0.30238 | 0.32773 | 0.3787 | 0.44121 | 0.51594
| High | Default | 0.15111 | 0.19148 | 0.21091 | 0.24367 | 0.28187 | 0.33191 | 0.38109 | 0.46046 | 0.35252 | 0.38192 | 0.43984 | 0.51231 | 0.60166
EP Mapping | Low | Default | 0.02952 | 0.037351 | 0.041196 | 0.046306 | 0.052559 | 0.059975 | 0.069877 | 0.077105 | 0.066218 | 0.072097 | 0.083917 | 0.098578 | 0.11581
| Norm | Default | 0.064964 | 0.081116 | 0.088901 | 0.10152 | 0.1215 | 0.1364 | 0.16561 | 0.18748 | 0.15104 | 0.16403 | 0.18926 | 0.2206 | 0.25807
| High | Default | 0.075502 | 0.095744 | 0.10542 | 0.12173 | 0.14103 | 0.16613 | 0.19065 | 0.23027 | 0.17643 | 0.19108 | 0.21986 | 0.25623 | 0.30113
Head | Low | Default | 0.13996 | 0.17485 | 0.19225 | 0.22053 | 0.25418 | 0.28784 | 0.32077 | 0.35401 | 0.29311 | 0.40231 | 0.45637 | 0.45637 | 0.45637
| Norm | Default | 0.28528 | 0.35312 | 0.38581 | 0.43946 | 0.51815 | 0.60182 | 0.66693 | 0.73685 | 0.63322 | 0.6781 | 0.761 | 0.86128 | 0.97624
| High | Default | 0.46196 | 0.56651 | 0.61773 | 0.69959 | 0.79601 | 0.92081 | 1.0404 | 1.2026 | 1.0004 | 1.0797 | 1.2365 | 1.4244 | 1.6435
Spine | Low | Default | 0.13996 | 0.17485 | 0.19225 | 0.22053 | 0.25418 | 0.28784 | 0.32077 | 0.35401 | 0.29311 | 0.40231 | 0.45637 | 0.45637 | 0.45637
| Norm | Default | 0.28528 | 0.35312 | 0.38581 | 0.43946 | 0.51815 | 0.60182 | 0.66693 | 0.73685 | 0.63322 | 0.6781 | 0.761 | 0.86128 | 0.97624
| High | Default | 0.46196 | 0.56651 | 0.61773 | 0.69959 | 0.79601 | 0.92081 | 1.0404 | 1.2026 | 1.0004 | 1.0797 | 1.2365 | 1.4244 | 1.6435
Thorax | Low | Default | 0.13996 | 0.17485 | 0.19225 | 0.22053 | 0.25418 | 0.28784 | 0.32077 | 0.35401 | 0.29311 | 0.40231 | 0.45637 | 0.45637 | 0.45637
| Norm | Default | 0.28528 | 0.35312 | 0.38581 | 0.43946 | 0.51815 | 0.60182 | 0.66693 | 0.73685 | 0.63322 | 0.6781 | 0.761 | 0.86128 | 0.97624
| High | Default | 0.46196 | 0.56651 | 0.61773 | 0.69959 | 0.79601 | 0.92081 | 1.0404 | 1.2026 | 1.0004 | 1.0797 | 1.2365 | 1.4244 | 1.6435
Abdomen | Low | Default | 0.13996 | 0.17485 | 0.19225 | 0.22053 | 0.25418 | 0.28784 | 0.32077 | 0.35401 | 0.29311 | 0.40231 | 0.45637 | 0.45637 | 0.45637
| Norm | Default | 0.28528 | 0.35312 | 0.38581 | 0.43946 | 0.51815 | 0.60182 | 0.66693 | 0.73685 | 0.63322 | 0.6781 | 0.761 | 0.86128 | 0.97624
| High | Default | 0.46196 | 0.56651 | 0.61773 | 0.69959 | 0.79601 | 0.92081 | 1.0404 | 1.2026 | 1.0004 | 1.0797 | 1.2365 | 1.4244 | 1.6435
Peripheral | Low | Default | 0.13996 | 0.17485 | 0.19225 | 0.22053 | 0.25418 | 0.28784 | 0.32077 | 0.35401 | 0.29311 | 0.40231 | 0.45637 | 0.45637 | 0.45637
| Norm | Default | 0.28528 | 0.35312 | 0.38581 | 0.43946 | 0.51815 | 0.60182 | 0.66693 | 0.73685 | 0.63322 | 0.6781 | 0.761 | 0.86128 | 0.97624
| High | Default | 0.46196 | 0.56651 | 0.61773 | 0.69959 | 0.79601 | 0.92081 | 1.0404 | 1.2026 | 1.0004 | 1.0797 | 1.2365 | 1.4244 | 1.6435
Radiography | Exposure | typical and worst case samples per application group
|  | field of View | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 0 | 1 | 2 | 3 | 4
|  | Field Size (cm): | 48 | 42 | 37 | 31 | 27 | 22 | 19 | 15 | 30 | 27 | 22 | 19 | 15
| X-ray Protocol | Patient Type | 0: Ref. AK (mGy/image) Frontal | Ref. AK (mGy/image) Lateral
Cardio | Left Coronary | Default | 0.060946 | 0.075785 | 0.083003 | 0.094646 | 0.10811 | 0.1254 | 0.14223 | 0.16831 | 0.13745 | 0.14841 | 0.1702 | 0.19672 | 0.22679
| 15 fps Low | Infant | 0.045977 | 0.054172 | 0.057997 | 0.064235 | 0.071418 | 0.080833 | 0.089966 | 0.10447 | 0.094614 | 0.10084 | 0.11297 | 0.12774 | 0.14535
|  | Large Adult, VeryLarge Adult | 0.060027 | 0.074583 | 0.081783 | 0.093188 | 0.10665 | 0.12376 | 0.14051 | 0.16624 | 0.13751 | 0.14847 | 0.17024 | 0.19648 | 0.22687
|  | Neonate | 0.025402 | 0.031708 | 0.034809 | 0.039999 | 0.046142 | 0.054275 | 0.062515 | 0.075798 | 0.062703 | 0.067853 | 0.078201 | 0.091181 | 0.10696
| Rotational Scan | Default | 0.11583 | 0.14345 | 0.15725 | 0.17874 | 0.20405 | 0.23628 | 0.26688 | 0.31443 | - | - | - | - | -
| Prop Ang0 4s | Infant | 0.039483 | 0.047742 | 0.05175 | 0.058256 | 0.065907 | 0.075919 | 0.085747 | 0.10143 | - | - | - | - | -
|  | Neonate | 0.025402 | 0.031708 | 0.034809 | 0.039999 | 0.046142 | 0.054275 | 0.062515 | 0.075798 | - | - | - | - | -
Pediatrics | 15 fps Contrast Normal | Default | 0.11797 | 0.14587 | 0.15961 | 0.18126 | 0.20607 | 0.23838 | 0.26888 | 0.31698 | 0.25171 | 0.27102 | 0.30998 | 0.3558 | 0.40759
|  | Child,SmallAdult | 0.069174 | 0.086047 | 0.093991 | 0.10731 | 0.12252 | 0.14231 | 0.16113 | 0.18823 | 0.1533 | 0.16578 | 0.18946 | 0.21926 | 0.25332
|  | Infant | 0.039483 | 0.047742 | 0.05175 | 0.058256 | 0.065907 | 0.075919 | 0.085747 | 0.10143 | 0.094614 | 0.10084 | 0.11297 | 0.12774 | 0.14535
|  | Neonate | 0.025402 | 0.031708 | 0.034809 | 0.039999 | 0.046142 | 0.054275 | 0.062515 | 0.075798 | 0.062703 | 0.067853 | 0.078201 | 0.091181 | 0.10696
| Prop Free Position | Default | 0.11409 | 0.14183 | 0.15495 | 0.17654 | 0.20167 | 0.2344 | 0.26482 | 0.3139 | - | - | - | - | -
|  | Neonate,Infant | 0.027337 | 0.033318 | 0.036249 | 0.041044 | 0.04669 | 0.054111 | 0.061582 | 0.073595 | - | - | - | - | -
EP | 15 Fps | Default | 0.069174 | 0.086047 | 0.093991 | 0.10731 | 0.12252 | 0.14231 | 0.16113 | 0.18823 | 0.1533 | 0.16578 | 0.18946 | 0.21926 | 0.25332
| Prop 4s | Default | 0.14139 | 0.18019 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | - | - | - | - | -
|  | Neonate,Infant | 0.09374 | 0.1207 | 0.15296 | 0.15296 | 0.15296 | 0.15296 | 0.15296 | 0.15296 | - | - | - | - | -
|  | LargeAdult,VeryLargeAdult | 0.14139 | 0.18019 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | - | - | - | - | -
EP Mapping | 7.5 fps low | Default | 0.010951 | 0.013772 | 0.015135 | 0.016971 | 0.01924 | 0.021877 | 0.025363 | 0.027939 | 0.024381 | 0.026573 | 0.030854 | 0.036146 | 0.042327
| 3D EP Prop 4s | Default | 0.14139 | 0.18019 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | - | - | - | - | -
|  | Neonate,Infant | 0.09374 | 0.1207 | 0.15296 | 0.15296 | 0.15296 | 0.15296 | 0.15296 | 0.15296 | - | - | - | - | -
|  | LargeAdult,VeryLargeAdult | 0.14139 | 0.18019 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | - | - | - | - | -
Head | Cerebral 2fps normal | Default | 2.8556 | 3.713 | 4.1392 | 4.8482 | 5.711 | 6.8524 | 7.8667 | 8.5486 | 7.227 | 7.5887 | 8.2764 | 9.0923 | 10.0487
|  | Child | 1.8389 | 2.2066 | 2.3841 | 2.677 | 3.0221 | 3.4817 | 3.9445 | 4.5343 | 4.2418 | 4.5496 | 5.1534 | 5.9039 | 6.8188
|  | Neonate,Infant | 1.8389 | 2.2066 | 2.3841 | 2.677 | 3.0221 | 3.4817 | 3.9445 | 4.5343 | 4.2418 | 4.5496 | 5.1534 | 5.9039 | 6.8188
| Aortic Arch 3 fps | Default | 4.2317 | 5.5019 | 6.1386 | 7.1975 | 8.1552 | 8.1552 | 8.1552 | 8.1552 | 10.5767 | 11.5808 | 12.2442 | 13.0061 | 13.8607
Spine | 4 fps | Default | 1.1966 | 1.5552 | 1.7342 | 2.0345 | 2.3904 | 2.8696 | 3.35 | 4.1265 | 2.4407 | 2.6841 | 3.1733 | 3.7828 | 4.5187
| 2 fps | Default | 1.4946 | 1.9442 | 2.1681 | 2.5404 | 2.9933 | 3.5911 | 4.1841 | 5.163 | 3.0499 | 3.3525 | 3.9619 | 4.7294 | 5.6457
Thorax | Lungs 2 fps | Default | 1.4272 | 1.857 | 2.0679 | 2.4255 | 2.6768 | 2.9379 | 3.1892 | 3.4919 | 3.6918 | 3.9149 | 4.2637 | 4.6827 | 5.1614
| Subclavian 3 fps | Default | 5.1357 | 6.6688 | 7.4467 | 8.7181 | 10.268 | 11.1833 | 11.737 | 12.5748 | 12.8193 | 14.9793 | 14.9793 | 14.9793 | 14.9793
Abdomen | Abdomen | Default | 2.6194 | 3.4057 | 3.8012 | 4.4551 | 5.2426 | 6.2961 | 7.3342 | 7.3342 | 5.6557 | 6.215 | 7.3512 | 8.7673 | 10.4564
| 3 fps | LargeAdult, VeryLargeAdult | 2.3063 | 2.995 | 3.3419 | 3.9224 | 4.6135 | 5.5302 | 6.4568 | 7.9533 | 4.985 | 5.4855 | 6.4753 | 7.7256 | 9.2172
| Iliac / Pelvis 3fps | Default | 4.2618 | 5.5375 | 6.1725 | 7.2433 | 8.5167 | 10.2184 | 11.9179 | 14.0411 | 10.8025 | 11.8816 | 14.0285 | 16.732 | 19.9563
Peripheral | Upper Legs 3 fps | Default | 3.6897 | 4.793 | 5.2918 | 5.6905 | 6.1423 | 6.7055 | 7.2411 | 7.8689 | 6.7619 | 7.0962 | 7.7495 | 8.5281 | 9.4327
| Lower Legs 1 fps | Default | 4.5091 | 5.62 | 5.9225 | 6.3608 | 6.8497 | 7.4536 | 8.0293 | 8.7079 | 7.2542 | 7.6078 | 8.2917 | 9.0906 | 10.0241
Roadmap |  | All modes for all application groups
|  | field of View | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 0 | 1 | 2 | 3 | 4
|  | Field Size (cm): | 48 | 42 | 37 | 31 | 27 | 22 | 19 | 15 | 30 | 27 | 22 | 19 | 15
| mode | Patient Type | 0: Ref. AK (mGy/image) Frontal | Ref. AK (mGy/image) Lateral
Cardio | Navigate | Default | 0.23267 | 0.28691 | 0.31309 | 0.35561 | 0.40488 | 0.46602 | 0.4959 | 0.4959 | 0.49287 | 0.62254 | 0.62254 | 0.62254 | 0.62254
| UnSubtract | Default | 0.23267 | 0.28691 | 0.31309 | 0.35561 | 0.40488 | 0.46602 | 0.4959 | 0.4959 | 0.49287 | 0.62254 | 0.62254 | 0.62254 | 0.62254
Pediatrics | Navigate | Default | 0.23267 | 0.28691 | 0.31309 | 0.35561 | 0.40488 | 0.46602 | 0.4959 | 0.4959 | 0.49287 | 0.62254 | 0.62254 | 0.62254 | 0.62254
| UnSubtract | Default | 0.23267 | 0.28691 | 0.31309 | 0.35561 | 0.40488 | 0.46602 | 0.4959 | 0.4959 | 0.49287 | 0.62254 | 0.62254 | 0.62254 | 0.62254
EP | Navigate | Default | 0.23267 | 0.28691 | 0.31309 | 0.35561 | 0.40488 | 0.46602 | 0.4959 | 0.4959 | 0.49287 | 0.62254 | 0.62254 | 0.62254 | 0.62254
EP Mapping | Navigate | Default | 0.23267 | 0.28691 | 0.31309 | 0.35561 | 0.40488 | 0.46602 | 0.4959 | 0.4959 | 0.49287 | 0.62254 | 0.62254 | 0.62254 | 0.62254
Head | Navigate | Default | 0.23267 | 0.28691 | 0.31309 | 0.35561 | 0.40488 | 0.46602 | 0.4959 | 0.4959 | 0.49287 | 0.62254 | 0.62254 | 0.62254 | 0.62254
| Carotid | Default | 0.20096 | 0.23961 | 0.25767 | 0.28687 | 0.32038 | 0.36306 | 0.38328 | 0.38328 | 0.46351 | 0.56688 | 0.56688 | 0.56688 | 0.56688
| Coil | Default | 1.0646 | 1.383 | 1.5438 | 1.8101 | 2.1291 | 2.5529 | 2.9837 | 3.6705 | 2.6907 | 2.9633 | 3.5006 | 4.1719 | 4.9826
Spine | UnSubtract | Default | 0.23267 | 0.28691 | 0.31309 | 0.35561 | 0.40488 | 0.46602 | 0.4959 | 0.4959 | 0.49287 | 0.62254 | 0.62254 | 0.62254 | 0.62254
| Navigate | Default | 0.23267 | 0.28691 | 0.31309 | 0.35561 | 0.40488 | 0.46602 | 0.4959 | 0.4959 | 0.49287 | 0.62254 | 0.62254 | 0.62254 | 0.62254
| Stent | Default | 0.20096 | 0.23961 | 0.25767 | 0.28687 | 0.32038 | 0.36306 | 0.38328 | 0.38328 | 0.46351 | 0.56688 | 0.56688 | 0.56688 | 0.56688
Thorax | UnSubtract | Default | 0.23267 | 0.28691 | 0.31309 | 0.35561 | 0.40488 | 0.46602 | 0.4959 | 0.4959 | 0.49287 | 0.62254 | 0.62254 | 0.62254 | 0.62254
| Navigate | Default | 0.23267 | 0.28691 | 0.31309 | 0.35561 | 0.40488 | 0.46602 | 0.4959 | 0.4959 | 0.49287 | 0.62254 | 0.62254 | 0.62254 | 0.62254
| Intervention | Default | 0.20096 | 0.23961 | 0.25767 | 0.28687 | 0.32038 | 0.36306 | 0.38328 | 0.38328 | 0.46351 | 0.56688 | 0.56688 | 0.56688 | 0.56688
Abdomen | UnSubtract | Default | 0.23267 | 0.28691 | 0.31309 | 0.35561 | 0.40488 | 0.46602 | 0.4959 | 0.4959 | 0.49287 | 0.62254 | 0.62254 | 0.62254 | 0.62254
| Navigate | Default | 0.23267 | 0.28691 | 0.31309 | 0.35561 | 0.40488 | 0.46602 | 0.4959 | 0.4959 | 0.49287 | 0.62254 | 0.62254 | 0.62254 | 0.62254
| Stent | Default | 0.20096 | 0.23961 | 0.25767 | 0.28687 | 0.32038 | 0.36306 | 0.38328 | 0.38328 | 0.46351 | 0.56688 | 0.56688 | 0.56688 | 0.56688
Peripheral | UnSubtract | Default | 0.23267 | 0.28691 | 0.31309 | 0.35561 | 0.40488 | 0.46602 | 0.4959 | 0.4959 | 0.49287 | 0.62254 | 0.62254 | 0.62254 | 0.62254
| Navigate | Default | 0.23267 | 0.28691 | 0.31309 | 0.35561 | 0.40488 | 0.46602 | 0.4959 | 0.4959 | 0.49287 | 0.62254 | 0.62254 | 0.62254 | 0.62254

## FD2012 Clarity
ch 16.20.10 | FD2012 Systems | Clarity
ch 16.20.10.1 | RAK values | 200 mm PMMA | Large SID
| Radioscopy | (Fluoro) | All modes for all application groups
|  | Field of View | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 0 | 1 | 2 | 3 | 4
|  | Field Size (cm) | 48 | 42 | 37 | 31 | 27 | 22 | 19 | 15 | 30 | 27 | 22 | 19 | 15
| Flavor | Patient Type | All modes for all application groups 0: Ref. AK (mGy/s) Frontal | Ref. AK (mGy/s) Lateral
Cardio Clarity | Low | Default | 0.089466 | 0.11052 | 0.12087 | 0.1376 | 0.15728 | 0.1813 | 0.20465 | 0.24135 | 0.2001 | 0.21667 | 0.24867 | 0.28781 | 0.33349
| Medium | Default | 0.13375 | 0.16604 | 0.18176 | 0.20761 | 0.23792 | 0.27776 | 0.31585 | 0.37732 | 0.29961 | 0.32468 | 0.37401 | 0.4342 | 0.5049
| Normal | Default | 0.3134 | 0.38455 | 0.41896 | 0.47434 | 0.53737 | 0.6205 | 0.70109 | 0.82654 | 0.67328 | 0.72613 | 0.82693 | 0.95167 | 1.0946
Cardio Pediatrics Clarity <40 kg | Low | Default | 0.038535 | 0.047964 | 0.052505 | 0.059964 | 0.068731 | 0.080253 | 0.091339 | 0.10908 | 0.087421 | 0.094777 | 0.10931 | 0.12709 | 0.148
| Medium | Default | 0.05691 | 0.071112 | 0.078044 | 0.089419 | 0.10271 | 0.11997 | 0.13665 | 0.15888 | 0.13037 | 0.14106 | 0.16031 | 0.18327 | 0.21051
| Normal | Default | 0.086768 | 0.10822 | 0.11841 | 0.13542 | 0.15497 | 0.17948 | 0.20299 | 0.24023 | 0.19829 | 0.21467 | 0.24733 | 0.28674 | 0.33311
Cardio Pediatrics Clarity >40 kg | Low | Default | 0.089466 | 0.11052 | 0.12087 | 0.1376 | 0.15728 | 0.1813 | 0.20465 | 0.24135 | 0.2001 | 0.21667 | 0.24867 | 0.28781 | 0.33349
| Medium | Default | 0.13375 | 0.16604 | 0.18176 | 0.20761 | 0.23792 | 0.27776 | 0.31585 | 0.37732 | 0.29961 | 0.32468 | 0.37401 | 0.4342 | 0.5049
| Normal | Default | 0.3134 | 0.38455 | 0.41896 | 0.47434 | 0.53737 | 0.6205 | 0.70109 | 0.82654 | 0.67328 | 0.72613 | 0.82693 | 0.95167 | 1.0946
EP Clarity | Low | Default | 0.035038 | 0.035038 | 0.035038 | 0.039761 | 0.045513 | 0.052866 | 0.059091 | 0.068227 | 0.060987 | 0.065879 | 0.075553 | 0.086084 | 0.097252
| Medium | Default | 0.069571 | 0.069571 | 0.069571 | 0.078995 | 0.090414 | 0.10497 | 0.11725 | 0.13596 | 0.12425 | 0.13426 | 0.15396 | 0.17438 | 0.19774
| Normal | Default | 0.089534 | 0.089534 | 0.089534 | 0.10235 | 0.11812 | 0.13841 | 0.15768 | 0.18881 | 0.14849 | 0.16109 | 0.18599 | 0.21671 | 0.25254
EP Mapping Clarity | Low | Default | 0.035038 | 0.035038 | 0.035038 | 0.039761 | 0.045513 | 0.052866 | 0.059091 | 0.068227 | 0.060987 | 0.065879 | 0.075553 | 0.086084 | 0.097252
| Medium | Default | 0.069571 | 0.069571 | 0.069571 | 0.078995 | 0.090414 | 0.10497 | 0.11725 | 0.13596 | 0.12425 | 0.13426 | 0.15396 | 0.17438 | 0.19774
| Normal | Default | 0.089534 | 0.089534 | 0.089534 | 0.10235 | 0.11812 | 0.13841 | 0.15768 | 0.18881 | 0.14849 | 0.16109 | 0.18599 | 0.21671 | 0.25254
Head Clarity | Low | Default | 0.12345 | 0.14803 | 0.15969 | 0.17841 | 0.20776 | 0.24378 | 0.27251 | 0.30479 | 0.26658 | 0.28724 | 0.32741 | 0.3751 | 0.43176
| Medium | Default | 0.18661 | 0.23367 | 0.25643 | 0.29404 | 0.34607 | 0.40182 | 0.44518 | 0.49251 | 0.42924 | 0.46068 | 0.52234 | 0.5985 | 0.68669
| Normal | Default | 0.33003 | 0.40911 | 0.44737 | 0.50875 | 0.58026 | 0.67228 | 0.76245 | 0.90546 | 0.74438 | 0.80407 | 0.92119 | 1.062 | 1.2295
Spine Clarity | Low | Default | 0.12345 | 0.14803 | 0.15969 | 0.17841 | 0.20776 | 0.24378 | 0.27251 | 0.30479 | 0.26658 | 0.28724 | 0.32741 | 0.3751 | 0.43176
| Medium | Default | 0.18661 | 0.23367 | 0.25643 | 0.29404 | 0.34607 | 0.40182 | 0.44518 | 0.49251 | 0.42924 | 0.46068 | 0.52234 | 0.5985 | 0.68669
| Normal | Default | 0.33003 | 0.40911 | 0.44737 | 0.50875 | 0.58026 | 0.67228 | 0.76245 | 0.90546 | 0.74438 | 0.80407 | 0.92119 | 1.062 | 1.2295
Thorax Clarity | Low | Default | 0.12345 | 0.14803 | 0.15969 | 0.17841 | 0.20776 | 0.24378 | 0.27251 | 0.30479 | 0.26658 | 0.28724 | 0.32741 | 0.3751 | 0.43176
| Medium | Default | 0.18661 | 0.23367 | 0.25643 | 0.29404 | 0.34607 | 0.40182 | 0.44518 | 0.49251 | 0.42924 | 0.46068 | 0.52234 | 0.5985 | 0.68669
| Normal | Default | 0.33003 | 0.40911 | 0.44737 | 0.50875 | 0.58026 | 0.67228 | 0.76245 | 0.90546 | 0.74438 | 0.80407 | 0.92119 | 1.062 | 1.2295
Abdomen Clarity | Low | Default | 0.12345 | 0.14803 | 0.15969 | 0.17841 | 0.20776 | 0.24378 | 0.27251 | 0.30479 | 0.26658 | 0.28724 | 0.32741 | 0.3751 | 0.43176
| Medium | Default | 0.18661 | 0.23367 | 0.25643 | 0.29404 | 0.34607 | 0.40182 | 0.44518 | 0.49251 | 0.42924 | 0.46068 | 0.52234 | 0.5985 | 0.68669
| Normal | Default | 0.33003 | 0.40911 | 0.44737 | 0.50875 | 0.58026 | 0.67228 | 0.76245 | 0.90546 | 0.74438 | 0.80407 | 0.92119 | 1.062 | 1.2295
Peripheral Clarity | Low | Default | 0.12345 | 0.14803 | 0.15969 | 0.17841 | 0.20776 | 0.24378 | 0.27251 | 0.30479 | 0.26658 | 0.28724 | 0.32741 | 0.3751 | 0.43176
| Medium | Default | 0.18661 | 0.23367 | 0.25643 | 0.29404 | 0.34607 | 0.40182 | 0.44518 | 0.49251 | 0.42924 | 0.46068 | 0.52234 | 0.5985 | 0.68669
| Normal | Default | 0.33003 | 0.40911 | 0.44737 | 0.50875 | 0.58026 | 0.67228 | 0.76245 | 0.90546 | 0.74438 | 0.80407 | 0.92119 | 1.062 | 1.2295
Radiography | Exposure | typical and worst case samples per application group
|  | field of View | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 0 | 1 | 2 | 3 | 4
|  | Field Size (cm): | 48 | 42 | 37 | 31 | 27 | 22 | 19 | 15 | 30 | 27 | 22 | 19 | 15
| X-ray Protocol | Patient Type | 0: Ref. AK (mGy/image) Frontal | Ref. AK (mGy/image) Lateral
Cardio Clarity | Left Coronary 15 fps Low | Default | 0.021771 | 0.027289 | 0.029996 | 0.034399 | 0.039464 | 0.046187 | 0.05264 | 0.062691 | 0.050367 | 0.05467 | 0.062994 | 0.073388 | 0.085379
| Rotational Scan Prop Ang0 4s | Default | 0.11801 | 0.14599 | 0.15949 | 0.18114 | 0.20609 | 0.23861 | 0.26891 | 0.31733 | - | - | - | - | -
|  | Infant | 0.039483 | 0.047742 | 0.05175 | 0.058256 | 0.065907 | 0.075919 | 0.085747 | 0.10143 | - | - | - | - | -
|  | Neonate | 0.025402 | 0.031708 | 0.034809 | 0.039999 | 0.046142 | 0.054275 | 0.062515 | 0.075798 | - | - | - | - | -
Cardio Pediatrics Clarity <40 kg | 15 fps Contrast low | Default | 0.011788 | 0.014679 | 0.016035 | 0.018272 | 0.020901 | 0.024316 | 0.027478 | 0.032403 | 0.026591 | 0.028712 | 0.032767 | 0.037709 | 0.043579
| 15 fps Contrast normal | Default | 0.018173 | 0.022951 | 0.025279 | 0.029165 | 0.0337 | 0.039602 | 0.04535 | 0.05245 | 0.04337 | 0.047231 | 0.054128 | 0.06149 | 0.069508
Cardio Pediatrics Clarity >40 kg | 15 fps Contrast low | Default | 0.021771 | 0.027289 | 0.029996 | 0.034399 | 0.039464 | 0.046187 | 0.05264 | 0.062691 | 0.050367 | 0.05467 | 0.062994 | 0.073388 | 0.085379
|  | Neonate,Infant,Child | 0.011788 | 0.014679 | 0.016035 | 0.018272 | 0.020901 | 0.024316 | 0.027478 | 0.032403 | 0.026591 | 0.028712 | 0.032767 | 0.037709 | 0.043579
| 15 fps Contrast normal | Default | 0.046649 | 0.05806 | 0.063583 | 0.072736 | 0.083181 | 0.096668 | 0.10979 | 0.1304 | 0.10187 | 0.11034 | 0.12646 | 0.14609 | 0.16892
|  | Neonate,Infant,Child | 0.018173 | 0.022951 | 0.025279 | 0.029165 | 0.0337 | 0.039602 | 0.04535 | 0.05245 | 0.04337 | 0.047231 | 0.054128 | 0.06149 | 0.069508
EP Clarity | 15 Fps | Default | 0.021771 | 0.027289 | 0.029996 | 0.034399 | 0.039464 | 0.046187 | 0.05264 | 0.062691 | 0.050367 | 0.05467 | 0.062994 | 0.073388 | 0.085379
| Prop 4s | Default | 0.14139 | 0.18019 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | - | - | - | - | -
|  | Neonate,Infant | 0.09374 | 0.1207 | 0.15296 | 0.15296 | 0.15296 | 0.15296 | 0.15296 | 0.15296 | - | - | - | - | -
|  | LargeAdult,VeryLargeAdult | 0.14139 | 0.18019 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | - | - | - | - | -
EP Mapping Clarity | 7.5 fps | Default | 0.021788 | 0.02729 | 0.030019 | 0.034421 | 0.039481 | 0.046161 | 0.052657 | 0.062787 | 0.050384 | 0.054607 | 0.062988 | 0.073311 | 0.085389
| 3D EP Prop 4s | Default | 0.14139 | 0.18019 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | - | - | - | - | -
|  | Neonate,Infant | 0.09374 | 0.1207 | 0.15296 | 0.15296 | 0.15296 | 0.15296 | 0.15296 | 0.15296 | - | - | - | - | -
|  | LargeAdult,VeryLargeAdult | 0.14139 | 0.18019 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | - | - | - | - | -
Head Clarity | Cerebral 2fps low | Default | 0.56967 | 0.74147 | 0.82723 | 0.96937 | 1.141 | 1.3673 | 1.5951 | 1.8818 | 1.4556 | 1.6018 | 1.8941 | 2.2584 | 2.6956
| Aortic Arch 3 fps | Default | 0.81451 | 1.059 | 1.1811 | 1.3847 | 1.629 | 1.956 | 2.2803 | 2.6887 | 2.0975 | 2.3053 | 2.725 | 3.2471 | 3.8781
Spine Clarity | 4 fps | Default | 0.67245 | 0.87457 | 0.9749 | 1.1424 | 1.3445 | 1.6137 | 1.8828 | 2.2207 | 1.6586 | 1.8214 | 2.1537 | 2.569 | 3.0642
| 2 fps | Default | 0.67298 | 0.87481 | 0.97398 | 1.1419 | 1.3443 | 1.6122 | 1.8824 | 2.2206 | 1.658 | 1.8238 | 2.1559 | 2.5701 | 3.0639
Thorax Clarity | Lungs 2 fps | Default | 0.71308 | 0.92798 | 1.0358 | 1.2126 | 1.4285 | 1.7128 | 1.999 | 2.3549 | 1.8415 | 2.022 | 2.3927 | 2.8491 | 3.4067
| Subclavian 3 fps | Default | 0.96496 | 1.2543 | 1.3975 | 1.6385 | 1.9281 | 2.3133 | 2.699 | 3.1798 | 2.4774 | 2.724 | 3.2218 | 3.8387 | 4.5848
Abdomen Clarity | 3 fps low | Default | 0.74759 | 0.97249 | 1.0838 | 1.2722 | 1.4967 | 1.7931 | 2.0953 | 2.4652 | 1.9269 | 2.1188 | 2.5048 | 2.9878 | 3.5643
| Iliac / Pelvis 3fps | Default | 0.71436 | 0.927 | 1.0345 | 1.214 | 1.4279 | 1.7144 | 1.9981 | 2.3556 | 1.8402 | 2.0237 | 2.3918 | 2.8497 | 3.4067
Peripheral Clarity | Upper Legs 3 fps | Default | 0.6311 | 0.80173 | 0.86798 | 0.97332 | 1.0973 | 1.236 | 1.3663 | 1.5232 | 1.5892 | 1.6854 | 1.8693 | 2.0926 | 2.3478
| Lower Legs 1 fps | Default | 1.1272 | 1.4649 | 1.6331 | 1.9154 | 2.1229 | 2.384 | 2.6252 | 2.9114 | 2.8969 | 3.1808 | 3.7603 | 4.258 | 4.7426
Roadmap |  | All modes for all application groups
|  | field of View | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 0 | 1 | 2 | 3 | 4
|  | Field Size (cm): | 48 | 42 | 37 | 31 | 27 | 22 | 19 | 15 | 30 | 27 | 22 | 19 | 15
| mode | Patient Type | 0: Ref. AK (mGy/image) Frontal | Ref. AK (mGy/image) Lateral
Cardio Clarity | Navigate | Default | 0.093115 | 0.11265 | 0.12193 | 0.13711 | 0.15437 | 0.17721 | 0.19865 | 0.23165 | 0.22177 | 0.23752 | 0.26804 | 0.30507 | 0.34608
| UnSubtract | Default | 0.093115 | 0.11265 | 0.12193 | 0.13711 | 0.15437 | 0.17721 | 0.19865 | 0.23165 | 0.22177 | 0.23752 | 0.26804 | 0.30507 | 0.34608
Cardio Pediatrics Clarity <40 kg | Navigate | Default | 0.093115 | 0.11265 | 0.12193 | 0.13711 | 0.15437 | 0.17721 | 0.19865 | 0.23165 | 0.22177 | 0.23752 | 0.26804 | 0.30507 | 0.34608
| UnSubtract | Default | 0.093115 | 0.11265 | 0.12193 | 0.13711 | 0.15437 | 0.17721 | 0.19865 | 0.23165 | 0.22177 | 0.23752 | 0.26804 | 0.30507 | 0.34608
Cardio Pediatrics Clarity >40 kg | Navigate | Default | 0.093115 | 0.11265 | 0.12193 | 0.13711 | 0.15437 | 0.17721 | 0.19865 | 0.23165 | 0.22177 | 0.23752 | 0.26804 | 0.30507 | 0.34608
| UnSubtract | Default | 0.093115 | 0.11265 | 0.12193 | 0.13711 | 0.15437 | 0.17721 | 0.19865 | 0.23165 | 0.22177 | 0.23752 | 0.26804 | 0.30507 | 0.34608
EP Clarity | Navigate | Default | 0.093115 | 0.11265 | 0.12193 | 0.13711 | 0.15437 | 0.17721 | 0.19865 | 0.23165 | 0.22177 | 0.23752 | 0.26804 | 0.30507 | 0.34608
| UnSubtract | Default | 0.093115 | 0.11265 | 0.12193 | 0.13711 | 0.15437 | 0.17721 | 0.19865 | 0.23165 | 0.22177 | 0.23752 | 0.26804 | 0.30507 | 0.34608
EP Mapping Clarity | Navigate | Default | 0.093115 | 0.11265 | 0.12193 | 0.13711 | 0.15437 | 0.17721 | 0.19865 | 0.23165 | 0.22177 | 0.23752 | 0.26804 | 0.30507 | 0.34608
| UnSubtract | Default | 0.093115 | 0.11265 | 0.12193 | 0.13711 | 0.15437 | 0.17721 | 0.19865 | 0.23165 | 0.22177 | 0.23752 | 0.26804 | 0.30507 | 0.34608
Head Clarity | Navigate | Default | 0.093115 | 0.11265 | 0.12193 | 0.13711 | 0.15437 | 0.17721 | 0.19865 | 0.23165 | 0.22177 | 0.23752 | 0.26804 | 0.30507 | 0.34608
| Coil | Default | 0.19155 | 0.24863 | 0.27769 | 0.32541 | 0.38291 | 0.45914 | 0.53579 | 0.66008 | 0.49314 | 0.54214 | 0.64086 | 0.76368 | 0.88972
| UnSubtract | Default | 0.093115 | 0.11265 | 0.12193 | 0.13711 | 0.15437 | 0.17721 | 0.19865 | 0.23165 | 0.22177 | 0.23752 | 0.26804 | 0.30507 | 0.34608
Spine Clarity | Navigate | Default | 0.093115 | 0.11265 | 0.12193 | 0.13711 | 0.15437 | 0.17721 | 0.19865 | 0.23165 | 0.22177 | 0.23752 | 0.26804 | 0.30507 | 0.34608
| Coil | Default | 0.19155 | 0.24863 | 0.27769 | 0.32541 | 0.38291 | 0.45914 | 0.53579 | 0.66008 | 0.49314 | 0.54214 | 0.64086 | 0.76368 | 0.88972
| UnSubtract | Default | 0.093115 | 0.11265 | 0.12193 | 0.13711 | 0.15437 | 0.17721 | 0.19865 | 0.23165 | 0.22177 | 0.23752 | 0.26804 | 0.30507 | 0.34608
Thorax Clarity | UnSubtract | Default | 0.093115 | 0.11265 | 0.12193 | 0.13711 | 0.15437 | 0.17721 | 0.19865 | 0.23165 | 0.22177 | 0.23752 | 0.26804 | 0.30507 | 0.34608
| Navigate | Default | 0.093115 | 0.11265 | 0.12193 | 0.13711 | 0.15437 | 0.17721 | 0.19865 | 0.23165 | 0.22177 | 0.23752 | 0.26804 | 0.30507 | 0.34608
Abdomen Clarity | UnSubtract | Default | 0.093115 | 0.11265 | 0.12193 | 0.13711 | 0.15437 | 0.17721 | 0.19865 | 0.23165 | 0.22177 | 0.23752 | 0.26804 | 0.30507 | 0.34608
| Navigate | Default | 0.093115 | 0.11265 | 0.12193 | 0.13711 | 0.15437 | 0.17721 | 0.19865 | 0.23165 | 0.22177 | 0.23752 | 0.26804 | 0.30507 | 0.34608
Peripheral Clarity | Navigate | Default | 0.093115 | 0.11265 | 0.12193 | 0.13711 | 0.15437 | 0.17721 | 0.19865 | 0.23165 | 0.22177 | 0.23752 | 0.26804 | 0.30507 | 0.34608
| UnSubtract | Default | 0.093115 | 0.11265 | 0.12193 | 0.13711 | 0.15437 | 0.17721 | 0.19865 | 0.23165 | 0.22177 | 0.23752 | 0.26804 | 0.30507 | 0.34608

## FD2015 Clarity
ch 16.20.12 | FD2015 Systems | Clarity
ch 16.20.12.1 | RAK values | 200 mm PMMA | Large SID
| Radioscopy | (Fluoro) | All modes for all application groups
|  | Field of View | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 0 | 1 | 2 | 3 | 4 | 5 | 6
|  | Field Size (cm) | 48 | 42 | 37 | 31 | 27 | 22 | 19 | 15 | 39 | 37 | 31 | 27 | 22 | 19 | 15
| Flavor | Patient Type | All modes for all application groups 0: Ref. AK (mGy/s) Frontal | Ref. AK (mGy/s) Lateral
Cardio Clarity | Low | Default | 0.086406 | 0.10687 | 0.11694 | 0.13296 | 0.15244 | 0.177 | 0.2015 | 0.24003 | 0.14637 | 0.15841 | 0.18205 | 0.20511 | 0.2393 | 0.27835 | 0.32704
| Medium | Default | 0.12967 | 0.16114 | 0.17662 | 0.2016 | 0.23102 | 0.26951 | 0.30737 | 0.36761 | 0.21686 | 0.23516 | 0.27135 | 0.30626 | 0.35793 | 0.41754 | 0.49157
| Normal | Default | 0.30109 | 0.36935 | 0.40267 | 0.45599 | 0.51665 | 0.59636 | 0.67406 | 0.79584 | 0.49492 | 0.53394 | 0.60932 | 0.68299 | 0.79105 | 0.91142 | 1.0635
Cardio Pediatrics Clarity <40 kg | Low | Default | 0.037355 | 0.046489 | 0.050957 | 0.058298 | 0.066814 | 0.07801 | 0.088909 | 0.10645 | 0.062481 | 0.067787 | 0.078214 | 0.08833 | 0.10351 | 0.12078 | 0.14262
| Medium | Default | 0.055372 | 0.069211 | 0.075901 | 0.086964 | 0.10008 | 0.1165 | 0.1325 | 0.15777 | 0.093175 | 0.10121 | 0.11695 | 0.13187 | 0.15346 | 0.17791 | 0.20668
| Normal | Default | 0.08429 | 0.10505 | 0.11519 | 0.13174 | 0.15125 | 0.17642 | 0.20122 | 0.24004 | 0.14129 | 0.15316 | 0.17655 | 0.19939 | 0.23267 | 0.27062 | 0.31849
Cardio Pediatrics Clarity >40 kg | Low | Default | 0.086406 | 0.10687 | 0.11694 | 0.13296 | 0.15244 | 0.177 | 0.2015 | 0.24003 | 0.14637 | 0.15841 | 0.18205 | 0.20511 | 0.2393 | 0.27835 | 0.32704
| Medium | Default | 0.12967 | 0.16114 | 0.17662 | 0.2016 | 0.23102 | 0.26951 | 0.30737 | 0.36761 | 0.21686 | 0.23516 | 0.27135 | 0.30626 | 0.35793 | 0.41754 | 0.49157
| Normal | Default | 0.30109 | 0.36935 | 0.40267 | 0.45599 | 0.51665 | 0.59636 | 0.67406 | 0.79584 | 0.49492 | 0.53394 | 0.60932 | 0.68299 | 0.79105 | 0.91142 | 1.0635
EP Clarity | Low | Default | 0.030093 | 0.030093 | 0.030093 | 0.034227 | 0.039119 | 0.045462 | 0.051367 | 0.060151 | 0.039427 | 0.0427 | 0.049083 | 0.055261 | 0.064302 | 0.074437 | 0.086115
| Medium | Default | 0.060664 | 0.060664 | 0.060664 | 0.068938 | 0.078914 | 0.09173 | 0.10365 | 0.12075 | 0.078629 | 0.084999 | 0.097595 | 0.10993 | 0.12773 | 0.14794 | 0.1706
| Normal | Default | 0.07526 | 0.07526 | 0.07526 | 0.086214 | 0.099328 | 0.11659 | 0.13263 | 0.1591 | 0.090058 | 0.097798 | 0.11271 | 0.12768 | 0.14962 | 0.17475 | 0.20604
EP Mapping Clarity | Low | Default | 0.035038 | 0.035038 | 0.035038 | 0.039761 | 0.045513 | 0.052866 | 0.059091 | 0.068227 | 0.04589 | 0.04961 | 0.056957 | 0.064153 | 0.074562 | 0.085326 | 0.097514
| Medium | Default | 0.069571 | 0.069571 | 0.069571 | 0.078995 | 0.090414 | 0.10497 | 0.11725 | 0.13596 | 0.091244 | 0.098662 | 0.11332 | 0.12745 | 0.14814 | 0.1694 | 0.1942
| Normal | Default | 0.089534 | 0.089534 | 0.089534 | 0.10235 | 0.11812 | 0.13841 | 0.15768 | 0.18881 | 0.10774 | 0.11711 | 0.1351 | 0.15299 | 0.17908 | 0.20891 | 0.24649
Head Clarity | Low | Default | 0.12262 | 0.1483 | 0.15978 | 0.1786 | 0.20771 | 0.24374 | 0.27251 | 0.30524 | 0.20416 | 0.21968 | 0.25008 | 0.28035 | 0.32263 | 0.37079 | 0.4328
| Medium | Default | 0.1854 | 0.23209 | 0.25508 | 0.29268 | 0.34566 | 0.4017 | 0.44511 | 0.49285 | 0.31985 | 0.34733 | 0.39928 | 0.44443 | 0.51111 | 0.58694 | 0.68263
| Normal | Default | 0.31961 | 0.39517 | 0.43247 | 0.49104 | 0.56092 | 0.65079 | 0.73813 | 0.87523 | 0.53235 | 0.5751 | 0.65913 | 0.7413 | 0.86138 | 0.99706 | 1.1557
Spine Clarity | Low | Default | 0.12262 | 0.1483 | 0.15978 | 0.1786 | 0.20771 | 0.24374 | 0.27251 | 0.30524 | 0.20416 | 0.21968 | 0.25008 | 0.28035 | 0.32263 | 0.37079 | 0.4328
| Medium | Default | 0.1854 | 0.23209 | 0.25508 | 0.29268 | 0.34566 | 0.4017 | 0.44511 | 0.49285 | 0.31985 | 0.34733 | 0.39928 | 0.44443 | 0.51111 | 0.58694 | 0.68263
| Normal | Default | 0.31961 | 0.39517 | 0.43247 | 0.49104 | 0.56092 | 0.65079 | 0.73813 | 0.87523 | 0.53235 | 0.5751 | 0.65913 | 0.7413 | 0.86138 | 0.99706 | 1.1557
Thorax Clarity | Low | Default | 0.12262 | 0.1483 | 0.15978 | 0.1786 | 0.20771 | 0.24374 | 0.27251 | 0.30524 | 0.20416 | 0.21968 | 0.25008 | 0.28035 | 0.32263 | 0.37079 | 0.4328
| Medium | Default | 0.1854 | 0.23209 | 0.25508 | 0.29268 | 0.34566 | 0.4017 | 0.44511 | 0.49285 | 0.31985 | 0.34733 | 0.39928 | 0.44443 | 0.51111 | 0.58694 | 0.68263
| Normal | Default | 0.31961 | 0.39517 | 0.43247 | 0.49104 | 0.56092 | 0.65079 | 0.73813 | 0.87523 | 0.53235 | 0.5751 | 0.65913 | 0.7413 | 0.86138 | 0.99706 | 1.1557
Abdomen Clarity | Low | Default | 0.12262 | 0.1483 | 0.15978 | 0.1786 | 0.20771 | 0.24374 | 0.27251 | 0.30524 | 0.20416 | 0.21968 | 0.25008 | 0.28035 | 0.32263 | 0.37079 | 0.4328
| Medium | Default | 0.1854 | 0.23209 | 0.25508 | 0.29268 | 0.34566 | 0.4017 | 0.44511 | 0.49285 | 0.31985 | 0.34733 | 0.39928 | 0.44443 | 0.51111 | 0.58694 | 0.68263
| Normal | Default | 0.31961 | 0.39517 | 0.43247 | 0.49104 | 0.56092 | 0.65079 | 0.73813 | 0.87523 | 0.53235 | 0.5751 | 0.65913 | 0.7413 | 0.86138 | 0.99706 | 1.1557
Peripheral Clarity | Low | Default | 0.12262 | 0.1483 | 0.15978 | 0.1786 | 0.20771 | 0.24374 | 0.27251 | 0.30524 | 0.20416 | 0.21968 | 0.25008 | 0.28035 | 0.32263 | 0.37079 | 0.4328
| Medium | Default | 0.1854 | 0.23209 | 0.25508 | 0.29268 | 0.34566 | 0.4017 | 0.44511 | 0.49285 | 0.31985 | 0.34733 | 0.39928 | 0.44443 | 0.51111 | 0.58694 | 0.68263
| Normal | Default | 0.31961 | 0.39517 | 0.43247 | 0.49104 | 0.56092 | 0.65079 | 0.73813 | 0.87523 | 0.53235 | 0.5751 | 0.65913 | 0.7413 | 0.86138 | 0.99706 | 1.1557
Radiography | Exposure | typical and worst case samples per application group
|  | field of View | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 0 | 1 | 2 | 3 | 4 | 5 | 6
|  | Field Size (cm): | 48 | 42 | 37 | 31 | 27 | 22 | 19 | 15 | 39 | 37 | 31 | 27 | 22 | 19 | 15
| X-ray Protocol | Patient Type | 0: Ref. AK (mGy/image) Frontal | Ref. AK (mGy/image) Lateral
Cardio Clarity | Left Coronary 15 fps Low | Default | 0.021771 | 0.027289 | 0.029996 | 0.034399 | 0.039464 | 0.046187 | 0.05264 | 0.062691 | 0.037091 | 0.040337 | 0.046655 | 0.052773 | 0.061681 | 0.071876 | 0.084667
| Rotational Scan Prop Ang0 4s | Default | 0.11801 | 0.14599 | 0.15949 | 0.18114 | 0.20609 | 0.23861 | 0.26891 | 0.31733 | - | - | - | - | - | - | -
|  | Infant | 0.039483 | 0.047742 | 0.05175 | 0.058256 | 0.065907 | 0.075919 | 0.085747 | 0.10143 | - | - | - | - | - | - | -
|  | Neonate | 0.025402 | 0.031708 | 0.034809 | 0.039999 | 0.046142 | 0.054275 | 0.062515 | 0.075798 | - | - | - | - | - | - | -
Cardio Pediatrics Clarity <40 kg | 15 fps Contrast low | Default | 0.011788 | 0.014679 | 0.016035 | 0.018272 | 0.020901 | 0.024316 | 0.027478 | 0.032403 | 0.02016 | 0.021847 | 0.025107 | 0.02833 | 0.032744 | 0.037855 | 0.044364
| 15 fps Contrast normal | Default | 0.018173 | 0.022951 | 0.025279 | 0.029165 | 0.0337 | 0.039602 | 0.04535 | 0.05245 | 0.031905 | 0.034763 | 0.040312 | 0.045832 | 0.053484 | 0.06094 | 0.069747
Cardio Pediatrics Clarity >40 kg | 15 fps Contrast low | Default | 0.021771 | 0.027289 | 0.029996 | 0.034399 | 0.039464 | 0.046187 | 0.05264 | 0.062691 | 0.037091 | 0.040337 | 0.046655 | 0.052773 | 0.061681 | 0.071876 | 0.084667
|  | Neonate,Infant,Child | 0.011788 | 0.014679 | 0.016035 | 0.018272 | 0.020901 | 0.024316 | 0.027478 | 0.032403 | 0.02016 | 0.021847 | 0.025107 | 0.02833 | 0.032744 | 0.037855 | 0.044364
| 15 fps Contrast normal | Default | 0.046649 | 0.05806 | 0.063583 | 0.072736 | 0.083181 | 0.096668 | 0.10979 | 0.1304 | 0.075175 | 0.081294 | 0.093543 | 0.10534 | 0.12221 | 0.14172 | 0.16571
|  | Neonate,Infant,Child | 0.018173 | 0.022951 | 0.025279 | 0.029165 | 0.0337 | 0.039602 | 0.04535 | 0.05245 | 0.031905 | 0.034763 | 0.040312 | 0.045832 | 0.053484 | 0.06094 | 0.069747
EP Clarity | 15 Fps | Default | 0.021771 | 0.027289 | 0.029996 | 0.034399 | 0.039464 | 0.046187 | 0.05264 | 0.062691 | 0.037091 | 0.040337 | 0.046655 | 0.052773 | 0.061681 | 0.071876 | 0.084667
| Prop 4s | Default | 0.14139 | 0.18019 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | - | - | - | - | - | - | -
|  | Neonate,Infant | 0.09374 | 0.1207 | 0.15296 | 0.15296 | 0.15296 | 0.15296 | 0.15296 | 0.15296 | - | - | - | - | - | - | -
|  | LargeAdult,VeryLargeAdult | 0.14139 | 0.18019 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | - | - | - | - | - | - | -
EP Mapping Clarity | 7.5 fps | Default | 0.021788 | 0.02729 | 0.030019 | 0.034421 | 0.039481 | 0.046161 | 0.052657 | 0.062787 | 0.037116 | 0.040345 | 0.046624 | 0.052771 | 0.061618 | 0.071808 | 0.084769
| 3D EP Prop 4s | Default | 0.14139 | 0.18019 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | - | - | - | - | - | - | -
|  | Neonate,Infant | 0.09374 | 0.1207 | 0.15296 | 0.15296 | 0.15296 | 0.15296 | 0.15296 | 0.15296 | - | - | - | - | - | - | -
|  | LargeAdult,VeryLargeAdult | 0.14139 | 0.18019 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | - | - | - | - | - | - | -
Head Clarity | Cerebral 2fps low | Default | 0.56967 | 0.74147 | 0.82723 | 0.96937 | 1.141 | 1.3673 | 1.5951 | 1.8818 | 1.0166 | 1.1189 | 1.3212 | 1.5238 | 1.8303 | 2.1842 | 2.6431
| Aortic Arch 3 fps | Default | 0.81451 | 1.059 | 1.1811 | 1.3847 | 1.629 | 1.956 | 2.2803 | 2.6887 | 1.4514 | 1.5967 | 1.8856 | 2.1755 | 2.6113 | 3.1198 | 3.771
Spine Clarity | 4 fps | Default | 0.67245 | 0.87457 | 0.9749 | 1.1424 | 1.3445 | 1.6137 | 1.8828 | 2.2207 | 1.1572 | 1.272 | 1.5045 | 1.7365 | 2.0838 | 2.4895 | 3.0075
| 2 fps | Default | 0.67298 | 0.87481 | 0.97398 | 1.1419 | 1.3443 | 1.6122 | 1.8824 | 2.2206 | 1.1575 | 1.272 | 1.5035 | 1.7357 | 2.0825 | 2.488 | 3.011
Thorax Clarity | Lungs 2 fps | Default | 0.71308 | 0.92798 | 1.0358 | 1.2126 | 1.4285 | 1.7128 | 1.999 | 2.3549 | 1.2807 | 1.4063 | 1.6635 | 1.9183 | 2.3031 | 2.7535 | 3.3271
| Subclavian 3 fps | Default | 0.96496 | 1.2543 | 1.3975 | 1.6385 | 1.9281 | 2.3133 | 2.699 | 3.1798 | 1.7059 | 1.8748 | 2.2169 | 2.5574 | 3.0659 | 3.6664 | 4.4336
Abdomen Clarity | 3 fps low | Default | 0.45402 | 0.59045 | 0.6586 | 0.77161 | 0.90734 | 1.0908 | 1.2716 | 1.4972 | 0.806 | 0.88663 | 1.0489 | 1.2102 | 1.4516 | 1.7325 | 2.0965
| Iliac / Pelvis 3fps | Default | 0.71436 | 0.927 | 1.0345 | 1.214 | 1.4279 | 1.7144 | 1.9981 | 2.3556 | 1.2806 | 1.4062 | 1.6639 | 1.9189 | 2.3027 | 2.7504 | 3.3236
Peripheral Clarity | Upper Legs 3 fps | Default | 0.6311 | 0.80173 | 0.86798 | 0.97332 | 1.0973 | 1.236 | 1.3663 | 1.5232 | 1.1308 | 1.2452 | 1.4705 | 1.6238 | 1.822 | 2.0433 | 2.3171
| Lower Legs 1 fps | Default | 1.1272 | 1.4649 | 1.6331 | 1.9154 | 2.1229 | 2.384 | 2.6252 | 2.9114 | 1.9794 | 2.1778 | 2.5775 | 2.9709 | 3.565 | 4.1331 | 4.6389
Roadmap |  | All modes for all application groups
|  | field of View | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 0 | 1 | 2 | 3 | 4 | 5 | 6
|  | Field Size (cm): | 48 | 42 | 37 | 31 | 27 | 22 | 19 | 15 | 39 | 37 | 31 | 27 | 22 | 19 | 15
| mode | Patient Type | 0: Ref. AK (mGy/image) Frontal | Ref. AK (mGy/image) Lateral
Cardio Clarity | Navigate | Default | 0.093115 | 0.11265 | 0.12193 | 0.13711 | 0.15437 | 0.17721 | 0.19865 | 0.23165 | 0.16954 | 0.18178 | 0.20517 | 0.22719 | 0.25975 | 0.29579 | 0.33999
| UnSubtract | Default | 0.093115 | 0.11265 | 0.12193 | 0.13711 | 0.15437 | 0.17721 | 0.19865 | 0.23165 | 0.16954 | 0.18178 | 0.20517 | 0.22719 | 0.25975 | 0.29579 | 0.33999
Cardio Pediatrics Clarity <40 kg | Navigate | Default | 0.093115 | 0.11265 | 0.12193 | 0.13711 | 0.15437 | 0.17721 | 0.19865 | 0.23165 | 0.16954 | 0.18178 | 0.20517 | 0.22719 | 0.25975 | 0.29579 | 0.33999
| UnSubtract | Default | 0.093115 | 0.11265 | 0.12193 | 0.13711 | 0.15437 | 0.17721 | 0.19865 | 0.23165 | 0.16954 | 0.18178 | 0.20517 | 0.22719 | 0.25975 | 0.29579 | 0.33999
Cardio Pediatrics Clarity >40 kg | Navigate | Default | 0.093115 | 0.11265 | 0.12193 | 0.13711 | 0.15437 | 0.17721 | 0.19865 | 0.23165 | 0.16954 | 0.18178 | 0.20517 | 0.22719 | 0.25975 | 0.29579 | 0.33999
| UnSubtract | Default | 0.093115 | 0.11265 | 0.12193 | 0.13711 | 0.15437 | 0.17721 | 0.19865 | 0.23165 | 0.16954 | 0.18178 | 0.20517 | 0.22719 | 0.25975 | 0.29579 | 0.33999
EP Clarity | Navigate | Default | 0.093115 | 0.11265 | 0.12193 | 0.13711 | 0.15437 | 0.17721 | 0.19865 | 0.23165 | 0.16954 | 0.18178 | 0.20517 | 0.22719 | 0.25975 | 0.29579 | 0.33999
| UnSubtract | Default | 0.093115 | 0.11265 | 0.12193 | 0.13711 | 0.15437 | 0.17721 | 0.19865 | 0.23165 | 0.16954 | 0.18178 | 0.20517 | 0.22719 | 0.25975 | 0.29579 | 0.33999
EP Mapping Clarity | Navigate | Default | 0.093115 | 0.11265 | 0.12193 | 0.13711 | 0.15437 | 0.17721 | 0.19865 | 0.23165 | 0.16954 | 0.18178 | 0.20517 | 0.22719 | 0.25975 | 0.29579 | 0.33999
| UnSubtract | Default | 0.093115 | 0.11265 | 0.12193 | 0.13711 | 0.15437 | 0.17721 | 0.19865 | 0.23165 | 0.16954 | 0.18178 | 0.20517 | 0.22719 | 0.25975 | 0.29579 | 0.33999
Head Clarity | Navigate | Default | 0.093115 | 0.11265 | 0.12193 | 0.13711 | 0.15437 | 0.17721 | 0.19865 | 0.23165 | 0.16954 | 0.18178 | 0.20517 | 0.22719 | 0.25975 | 0.29579 | 0.33999
| Coil | Default | 0.19155 | 0.24863 | 0.27769 | 0.32541 | 0.38291 | 0.45914 | 0.53579 | 0.66008 | 0.34055 | 0.37449 | 0.44268 | 0.51102 | 0.61351 | 0.73276 | 0.8723
| UnSubtract | Default | 0.093115 | 0.11265 | 0.12193 | 0.13711 | 0.15437 | 0.17721 | 0.19865 | 0.23165 | 0.16954 | 0.18178 | 0.20517 | 0.22719 | 0.25975 | 0.29579 | 0.33999
Spine Clarity | Navigate | Default | 0.093115 | 0.11265 | 0.12193 | 0.13711 | 0.15437 | 0.17721 | 0.19865 | 0.23165 | 0.16954 | 0.18178 | 0.20517 | 0.22719 | 0.25975 | 0.29579 | 0.33999
| Coil | Default | 0.19155 | 0.24863 | 0.27769 | 0.32541 | 0.38291 | 0.45914 | 0.53579 | 0.66008 | 0.34055 | 0.37449 | 0.44268 | 0.51102 | 0.61351 | 0.73276 | 0.8723
| UnSubtract | Default | 0.093115 | 0.11265 | 0.12193 | 0.13711 | 0.15437 | 0.17721 | 0.19865 | 0.23165 | 0.16954 | 0.18178 | 0.20517 | 0.22719 | 0.25975 | 0.29579 | 0.33999
Thorax Clarity | UnSubtract | Default | 0.093115 | 0.11265 | 0.12193 | 0.13711 | 0.15437 | 0.17721 | 0.19865 | 0.23165 | 0.16954 | 0.18178 | 0.20517 | 0.22719 | 0.25975 | 0.29579 | 0.33999
| Navigate | Default | 0.093115 | 0.11265 | 0.12193 | 0.13711 | 0.15437 | 0.17721 | 0.19865 | 0.23165 | 0.16954 | 0.18178 | 0.20517 | 0.22719 | 0.25975 | 0.29579 | 0.33999
Abdomen Clarity | UnSubtract | Default | 0.093115 | 0.11265 | 0.12193 | 0.13711 | 0.15437 | 0.17721 | 0.19865 | 0.23165 | 0.16954 | 0.18178 | 0.20517 | 0.22719 | 0.25975 | 0.29579 | 0.33999
| Navigate | Default | 0.093115 | 0.11265 | 0.12193 | 0.13711 | 0.15437 | 0.17721 | 0.19865 | 0.23165 | 0.16954 | 0.18178 | 0.20517 | 0.22719 | 0.25975 | 0.29579 | 0.33999
Peripheral Clarity | Navigate | Default | 0.093115 | 0.11265 | 0.12193 | 0.13711 | 0.15437 | 0.17721 | 0.19865 | 0.23165 | 0.16954 | 0.18178 | 0.20517 | 0.22719 | 0.25975 | 0.29579 | 0.33999
| UnSubtract | Default | 0.093115 | 0.11265 | 0.12193 | 0.13711 | 0.15437 | 0.17721 | 0.19865 | 0.23165 | 0.16954 | 0.18178 | 0.20517 | 0.22719 | 0.25975 | 0.29579 | 0.33999

## FD2015
ch 16.20.11 | FD2015 Systems
ch 16.20.11.1 | RAK values | 200 mm PMMA | Large SID
| Radioscopy | (Fluoro) | All modes for all application groups
|  | Field of View | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 0 | 1 | 2 | 3 | 4 | 5 | 6
|  | Field Size (cm) | 48 | 42 | 37 | 31 | 27 | 22 | 19 | 15 | 39 | 37 | 31 | 27 | 22 | 19 | 15
| Flavor | Patient Type | All modes for all application groups 0: Ref. AK (mGy/s) Frontal | Ref. AK (mGy/s) Lateral
Cardio | Low | Default | 0.1223 | 0.15481 | 0.17092 | 0.19712 | 0.22786 | 0.26863 | 0.30851 | 0.37239 | 0.26195 | 0.28494 | 0.3305 | 0.37562 | 0.44255 | 0.51866 | 0.61557
| Norm | Default | 0.32327 | 0.3934 | 0.42632 | 0.48017 | 0.54261 | 0.62294 | 0.6955 | 0.78576 | 0.52426 | 0.56382 | 0.64055 | 0.7136 | 0.82274 | 0.94073 | 1.0537
| High | Default | 0.42296 | 0.52058 | 0.56779 | 0.64292 | 0.73021 | 0.84219 | 0.95071 | 1.1117 | 0.69647 | 0.7505 | 0.85834 | 0.96088 | 1.1117 | 1.2816 | 1.4927
Pediatrics | Low | Default | 0.11218 | 0.14014 | 0.15378 | 0.1763 | 0.20262 | 0.23716 | 0.27094 | 0.30876 | 0.18808 | 0.20422 | 0.23595 | 0.26729 | 0.31297 | 0.36578 | 0.41783
| Norm | Default | 0.25304 | 0.31318 | 0.34261 | 0.39007 | 0.44262 | 0.50052 | 0.55692 | 0.64479 | 0.45072 | 0.48869 | 0.56133 | 0.6278 | 0.71526 | 0.81024 | 0.92963
| High | Default | 0.42759 | 0.51388 | 0.55313 | 0.61486 | 0.66286 | 0.7243 | 0.78171 | 0.87065 | 0.68004 | 0.72797 | 0.81783 | 0.90467 | 0.99286 | 1.0874 | 1.2017
EP | Low | Default | 0.05789 | 0.073345 | 0.080892 | 0.090916 | 0.10335 | 0.11785 | 0.13571 | 0.14907 | 0.094846 | 0.10358 | 0.12055 | 0.1372 | 0.16231 | 0.18989 | 0.22433
| Norm | Default | 0.12653 | 0.15756 | 0.17258 | 0.1975 | 0.23702 | 0.26637 | 0.32466 | 0.36799 | 0.2203 | 0.23892 | 0.27575 | 0.31257 | 0.36697 | 0.43057 | 0.51205
| High | Default | 0.14848 | 0.18774 | 0.20711 | 0.23909 | 0.27657 | 0.32597 | 0.37463 | 0.45228 | 0.27058 | 0.29348 | 0.33842 | 0.3826 | 0.448 | 0.52289 | 0.61724
EP Mapping | Low | Default | 0.02952 | 0.037351 | 0.041196 | 0.046306 | 0.052559 | 0.059975 | 0.069877 | 0.077105 | 0.048219 | 0.052589 | 0.061183 | 0.069805 | 0.082347 | 0.096856 | 0.11543
| Norm | Default | 0.064964 | 0.081116 | 0.088901 | 0.10152 | 0.1215 | 0.1364 | 0.16561 | 0.18748 | 0.11311 | 0.12267 | 0.14133 | 0.16009 | 0.18752 | 0.2192 | 0.25999
| High | Default | 0.075502 | 0.095744 | 0.10542 | 0.12173 | 0.14103 | 0.16613 | 0.19065 | 0.23027 | 0.13354 | 0.14547 | 0.16861 | 0.19149 | 0.22561 | 0.26439 | 0.31357
Head | Low | Default | 0.13597 | 0.17057 | 0.18761 | 0.2156 | 0.249 | 0.28193 | 0.31503 | 0.3483 | 0.21171 | 0.23139 | 0.27116 | 0.31058 | 0.36966 | 0.43843 | 0.5166
| Norm | Default | 0.27548 | 0.34098 | 0.37343 | 0.42621 | 0.50277 | 0.58918 | 0.65659 | 0.73072 | 0.49664 | 0.53142 | 0.60045 | 0.66812 | 0.76669 | 0.87805 | 1.0192
| High | Default | 0.4435 | 0.5451 | 0.59449 | 0.67357 | 0.76726 | 0.88653 | 0.98652 | 1.1224 | 0.77395 | 0.83531 | 0.95723 | 1.0752 | 1.2482 | 1.4437 | 1.6683
Spine | Low | Default | 0.13597 | 0.17057 | 0.18761 | 0.2156 | 0.249 | 0.28193 | 0.31503 | 0.3483 | 0.21171 | 0.23139 | 0.27116 | 0.31058 | 0.36966 | 0.43843 | 0.5166
| Norm | Default | 0.27548 | 0.34098 | 0.37343 | 0.42621 | 0.50277 | 0.58918 | 0.65659 | 0.73072 | 0.49664 | 0.53142 | 0.60045 | 0.66812 | 0.76669 | 0.87805 | 1.0192
| High | Default | 0.4435 | 0.5451 | 0.59449 | 0.67357 | 0.76726 | 0.88653 | 0.98652 | 1.1224 | 0.77395 | 0.83531 | 0.95723 | 1.0752 | 1.2482 | 1.4437 | 1.6683
Thorax | Low | Default | 0.13597 | 0.17057 | 0.18761 | 0.2156 | 0.249 | 0.28193 | 0.31503 | 0.3483 | 0.21171 | 0.23139 | 0.27116 | 0.31058 | 0.36966 | 0.43843 | 0.5166
| Norm | Default | 0.27548 | 0.34098 | 0.37343 | 0.42621 | 0.50277 | 0.58918 | 0.65659 | 0.73072 | 0.49664 | 0.53142 | 0.60045 | 0.66812 | 0.76669 | 0.87805 | 1.0192
| High | Default | 0.4435 | 0.5451 | 0.59449 | 0.67357 | 0.76726 | 0.88653 | 0.98652 | 1.1224 | 0.77395 | 0.83531 | 0.95723 | 1.0752 | 1.2482 | 1.4437 | 1.6683
Abdomen | Low | Default | 0.13597 | 0.17057 | 0.18761 | 0.2156 | 0.249 | 0.28193 | 0.31503 | 0.3483 | 0.21171 | 0.23139 | 0.27116 | 0.31058 | 0.36966 | 0.43843 | 0.5166
| Norm | Default | 0.27548 | 0.34098 | 0.37343 | 0.42621 | 0.50277 | 0.58918 | 0.65659 | 0.73072 | 0.49664 | 0.53142 | 0.60045 | 0.66812 | 0.76669 | 0.87805 | 1.0192
| High | Default | 0.4435 | 0.5451 | 0.59449 | 0.67357 | 0.76726 | 0.88653 | 0.98652 | 1.1224 | 0.77395 | 0.83531 | 0.95723 | 1.0752 | 1.2482 | 1.4437 | 1.6683
Peripheral | Low | Default | 0.13597 | 0.17057 | 0.18761 | 0.2156 | 0.249 | 0.28193 | 0.31503 | 0.3483 | 0.21171 | 0.23139 | 0.27116 | 0.31058 | 0.36966 | 0.43843 | 0.5166
| Norm | Default | 0.27548 | 0.34098 | 0.37343 | 0.42621 | 0.50277 | 0.58918 | 0.65659 | 0.73072 | 0.49664 | 0.53142 | 0.60045 | 0.66812 | 0.76669 | 0.87805 | 1.0192
| High | Default | 0.4435 | 0.5451 | 0.59449 | 0.67357 | 0.76726 | 0.88653 | 0.98652 | 1.1224 | 0.77395 | 0.83531 | 0.95723 | 1.0752 | 1.2482 | 1.4437 | 1.6683
Radiography | Exposure | typical and worst case samples per application group
|  | field of View | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 0 | 1 | 2 | 3 | 4 | 5 | 6
|  | Field Size (cm): | 48 | 42 | 37 | 31 | 27 | 22 | 19 | 15 | 39 | 37 | 31 | 27 | 22 | 19 | 15
| X-ray Protocol | Patient Type | 0: Ref. AK (mGy/image) Frontal | Ref. AK (mGy/image) Lateral
Cardio | Left Coronary | Default | 0.060946 | 0.075785 | 0.083003 | 0.094646 | 0.10811 | 0.1254 | 0.14223 | 0.16831 | 0.099692 | 0.10778 | 0.12388 | 0.13924 | 0.16202 | 0.1876 | 0.21887
| 15 fps Low | Infant | 0.045977 | 0.054172 | 0.057997 | 0.064235 | 0.071418 | 0.080833 | 0.089966 | 0.10447 | 0.069538 | 0.074148 | 0.082955 | 0.091694 | 0.10453 | 0.11924 | 0.13811
|  | Large Adult, VeryLarge Adult | 0.060027 | 0.074583 | 0.081783 | 0.093188 | 0.10665 | 0.12376 | 0.14051 | 0.16624 | 0.099676 | 0.10773 | 0.12394 | 0.13924 | 0.16193 | 0.18754 | 0.2189
|  | Neonate | 0.025402 | 0.031708 | 0.034809 | 0.039999 | 0.046142 | 0.054275 | 0.062515 | 0.075798 | 0.045618 | 0.049347 | 0.056924 | 0.064447 | 0.075879 | 0.089407 | 0.092668
| Rotational Scan | Default | 0.11583 | 0.14345 | 0.15725 | 0.17874 | 0.20405 | 0.23628 | 0.26688 | 0.31443 | - | - | - | - | - | - | -
| Prop Ang0 4s | Infant | 0.039483 | 0.047742 | 0.05175 | 0.058256 | 0.065907 | 0.075919 | 0.085747 | 0.10143 | - | - | - | - | - | - | -
|  | Neonate | 0.025402 | 0.031708 | 0.034809 | 0.039999 | 0.046142 | 0.054275 | 0.062515 | 0.075798 | - | - | - | - | - | - | -
Pediatrics | 15 fps Contrast Normal | Default | 0.11797 | 0.14587 | 0.15961 | 0.18126 | 0.20607 | 0.23838 | 0.26888 | 0.31698 | 0.18679 | 0.20183 | 0.23076 | 0.25838 | 0.29877 | 0.34398 | 0.39933
|  | Child,SmallAdult | 0.069174 | 0.086047 | 0.093991 | 0.10731 | 0.12252 | 0.14231 | 0.16113 | 0.18823 | 0.11316 | 0.12228 | 0.14058 | 0.15794 | 0.18354 | 0.21202 | 0.24809
|  | Infant | 0.039483 | 0.047742 | 0.05175 | 0.058256 | 0.065907 | 0.075919 | 0.085747 | 0.10143 | 0.066796 | 0.071721 | 0.081446 | 0.091052 | 0.10505 | 0.1212 | 0.14155
|  | Neonate | 0.025402 | 0.031708 | 0.034809 | 0.039999 | 0.046142 | 0.054275 | 0.062515 | 0.075798 | 0.0454 | 0.049379 | 0.057172 | 0.065054 | 0.076719 | 0.090475 | 0.10826
| Prop Free Position | Default | 0.11409 | 0.14183 | 0.15495 | 0.17654 | 0.20167 | 0.2344 | 0.26482 | 0.3139 | - | - | - | - | - | - | -
|  | Neonate,Infant | 0.027337 | 0.033318 | 0.036249 | 0.041044 | 0.04669 | 0.054111 | 0.061582 | 0.073595 | - | - | - | - | - | - | -
EP | 15 Fps | Default | 0.069174 | 0.086047 | 0.093991 | 0.10731 | 0.12252 | 0.14231 | 0.16113 | 0.18823 | 0.11316 | 0.12228 | 0.14058 | 0.15794 | 0.18354 | 0.21202 | 0.24809
| Prop 4s | Default | 0.14139 | 0.18019 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | - | - | - | - | - | - | -
|  | Neonate,Infant | 0.09374 | 0.1207 | 0.15296 | 0.15296 | 0.15296 | 0.15296 | 0.15296 | 0.15296 | - | - | - | - | - | - | -
|  | LargeAdult,VeryLargeAdult | 0.14139 | 0.18019 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | - | - | - | - | - | - | -
EP Mapping | 7.5 fps low | Default | 0.010951 | 0.013772 | 0.015135 | 0.016971 | 0.01924 | 0.021877 | 0.025363 | 0.027939 | 0.017835 | 0.019445 | 0.022566 | 0.02566 | 0.030227 | 0.035504 | 0.042205
| 3D EP Prop 4s | Default | 0.14139 | 0.18019 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | - | - | - | - | - | - | -
|  | Neonate,Infant | 0.09374 | 0.1207 | 0.15296 | 0.15296 | 0.15296 | 0.15296 | 0.15296 | 0.15296 | - | - | - | - | - | - | -
|  | LargeAdult,VeryLargeAdult | 0.14139 | 0.18019 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | 0.22513 | - | - | - | - | - | - | -
Head | Cerebral 2fps normal | Default | 2.8556 | 3.713 | 4.1392 | 4.8482 | 5.711 | 6.8524 | 7.8667 | 8.5486 | 5.0926 | 5.5966 | 6.6151 | 7.2581 | 7.9852 | 8.8072 | 9.8202
|  | Child | 1.8389 | 2.2066 | 2.3841 | 2.677 | 3.0221 | 3.4817 | 3.9445 | 4.5343 | 3.2532 | 3.4817 | 3.9253 | 4.3677 | 5.019 | 5.7876 | 6.7839
|  | Neonate,Infant | 1.8389 | 2.2066 | 2.3841 | 2.677 | 3.0221 | 3.4817 | 3.9445 | 4.5343 | 3.2532 | 3.4817 | 3.9253 | 4.3677 | 5.019 | 5.7876 | 6.7839
| Aortic Arch 3 fps | Default | 4.2317 | 5.5019 | 6.1386 | 7.1975 | 8.1552 | 8.1552 | 8.1552 | 8.1552 | 7.3506 | 8.0973 | 9.5605 | 11.0305 | 11.0305 | 11.0305 | 11.0305
Spine | 4 fps | Default | 1.1953 | 1.556 | 1.7337 | 2.0345 | 2.3927 | 2.8691 | 3.3524 | 4.1257 | 1.7046 | 1.8767 | 2.2159 | 2.5595 | 3.068 | 3.662 | 4.4359
| 2 fps | Default | 1.4946 | 1.9442 | 2.1681 | 2.5404 | 2.9933 | 3.5911 | 4.1841 | 5.163 | 2.1349 | 2.3461 | 2.77 | 3.1975 | 3.8374 | 4.5867 | 5.5504
Thorax | Lungs 2 fps | Default | 1.4275 | 1.8563 | 2.0707 | 2.4279 | 2.6574 | 2.9194 | 3.1692 | 3.4736 | 2.5607 | 2.8145 | 3.3275 | 3.7842 | 4.1588 | 4.5808 | 5.0978
| Subclavian 3 fps | Default | 5.1357 | 6.6688 | 7.4467 | 8.7181 | 10.268 | 11.1833 | 11.737 | 12.5748 | 8.8631 | 9.7504 | 11.5325 | 13.2857 | 14.4072 | 15.7565 | 15.7565
Abdomen | Abdomen | Default | 2.6194 | 3.4057 | 3.8012 | 4.4551 | 5.2426 | 6.2961 | 7.3342 | 7.3342 | 4.5969 | 5.056 | 5.9724 | 6.8961 | 8.2732 | 9.885 | 10.4746
| 3 fps | LargeAdult, VeryLargeAdult | 2.3063 | 2.995 | 3.3419 | 3.9224 | 4.6135 | 5.5302 | 6.4568 | 7.9533 | 4.0627 | 4.4741 | 5.2885 | 6.1054 | 7.326 | 8.7345 | 10.5677
| Iliac / Pelvis 3fps | Default | 4.2618 | 5.5375 | 6.1725 | 7.2433 | 8.5167 | 10.2184 | 11.9179 | 14.0411 | 7.4709 | 8.2093 | 9.7082 | 11.2135 | 13.4456 | 16.0441 | 19.4319
Peripheral | Upper Legs 3 fps | Default | 3.6897 | 4.793 | 5.2918 | 5.6905 | 6.1423 | 6.7055 | 7.2411 | 7.8689 | 5.5588 | 5.8226 | 6.331 | 6.821 | 7.5137 | 8.299 | 9.2678
| Lower Legs 1 fps | Default | 4.5091 | 5.62 | 5.9225 | 6.3608 | 6.8497 | 7.4536 | 8.0293 | 8.7079 | 5.969 | 6.2469 | 6.7777 | 7.2863 | 8.0129 | 8.8101 | 9.8087
Roadmap |  | All modes for all application groups
|  | field of View | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 0 | 1 | 2 | 3 | 4 | 5 | 6
|  | Field Size (cm): | 48 | 42 | 37 | 31 | 27 | 22 | 19 | 15 | 39 | 37 | 31 | 27 | 22 | 19 | 15
| mode | Patient Type | 0: Ref. AK (mGy/image) Frontal | Ref. AK (mGy/image) Lateral
Cardio | Navigate | Default | 0.23267 | 0.28691 | 0.31309 | 0.35561 | 0.40488 | 0.46602 | 0.4959 | 0.4959 | 0.36687 | 0.39585 | 0.45221 | 0.47355 | 0.57848 | 0.67088 | 0.77768
| UnSubtract | Default | 0.23267 | 0.28691 | 0.31309 | 0.35561 | 0.40488 | 0.46602 | 0.4959 | 0.4959 | 0.36687 | 0.39585 | 0.45221 | 0.47355 | 0.57848 | 0.67088 | 0.77768
Pediatrics | Navigate | Default | 0.23267 | 0.28691 | 0.31309 | 0.35561 | 0.40488 | 0.46602 | 0.4959 | 0.4959 | 0.36687 | 0.39585 | 0.45221 | 0.47355 | 0.57848 | 0.67088 | 0.77768
| UnSubtract | Default | 0.23267 | 0.28691 | 0.31309 | 0.35561 | 0.40488 | 0.46602 | 0.4959 | 0.4959 | 0.36687 | 0.39585 | 0.45221 | 0.47355 | 0.57848 | 0.67088 | 0.77768
EP | Navigate | Default | 0.23267 | 0.28691 | 0.31309 | 0.35561 | 0.40488 | 0.46602 | 0.4959 | 0.4959 | 0.36687 | 0.39585 | 0.45221 | 0.47355 | 0.57848 | 0.67088 | 0.77768
EP Mapping | Navigate | Default | 0.23267 | 0.28691 | 0.31309 | 0.35561 | 0.40488 | 0.46602 | 0.4959 | 0.4959 | 0.36687 | 0.39585 | 0.45221 | 0.47355 | 0.57848 | 0.67088 | 0.77768
Head | Navigate | Default | 0.23267 | 0.28691 | 0.31309 | 0.35561 | 0.40488 | 0.46602 | 0.4959 | 0.4959 | 0.36495 | 0.39435 | 0.44987 | 0.47194 | 0.57671 | 0.62331 | 0.62331
| Carotid | Default | 0.20096 | 0.23961 | 0.25767 | 0.28687 | 0.32038 | 0.36306 | 0.38328 | 0.38328 | 0.3617 | 0.38467 | 0.43057 | 0.4734 | 0.53445 | 0.56793 | 0.56793
| Coil | Default | 1.0646 | 1.383 | 1.5438 | 1.8101 | 2.1291 | 2.5529 | 2.9837 | 3.6705 | 1.8655 | 2.0521 | 2.4289 | 2.7994 | 3.3614 | 4.0185 | 4.8522
Spine | UnSubtract | Default | 0.23267 | 0.28691 | 0.31309 | 0.35561 | 0.40488 | 0.46602 | 0.4959 | 0.4959 | 0.36495 | 0.39435 | 0.44987 | 0.47194 | 0.57671 | 0.62331 | 0.62331
| Navigate | Default | 0.23267 | 0.28691 | 0.31309 | 0.35561 | 0.40488 | 0.46602 | 0.4959 | 0.4959 | 0.36495 | 0.39435 | 0.44987 | 0.47194 | 0.57671 | 0.62331 | 0.62331
| Stent | Default | 0.20096 | 0.23961 | 0.25767 | 0.28687 | 0.32038 | 0.36306 | 0.38328 | 0.38328 | 0.3617 | 0.38467 | 0.43057 | 0.4734 | 0.53445 | 0.56793 | 0.56793
Thorax | UnSubtract | Default | 0.23267 | 0.28691 | 0.31309 | 0.35561 | 0.40488 | 0.46602 | 0.4959 | 0.4959 | 0.36495 | 0.39435 | 0.44987 | 0.47194 | 0.57671 | 0.62331 | 0.62331
| Navigate | Default | 0.23267 | 0.28691 | 0.31309 | 0.35561 | 0.40488 | 0.46602 | 0.4959 | 0.4959 | 0.36495 | 0.39435 | 0.44987 | 0.47194 | 0.57671 | 0.62331 | 0.62331
| Intervention | Default | 0.20096 | 0.23961 | 0.25767 | 0.28687 | 0.32038 | 0.36306 | 0.38328 | 0.38328 | 0.3617 | 0.38467 | 0.43057 | 0.4734 | 0.53445 | 0.56793 | 0.56793
Abdomen | UnSubtract | Default | 0.23267 | 0.28691 | 0.31309 | 0.35561 | 0.40488 | 0.46602 | 0.4959 | 0.4959 | 0.36495 | 0.39435 | 0.44987 | 0.47194 | 0.57671 | 0.62331 | 0.62331
| Navigate | Default | 0.23267 | 0.28691 | 0.31309 | 0.35561 | 0.40488 | 0.46602 | 0.4959 | 0.4959 | 0.36495 | 0.39435 | 0.44987 | 0.47194 | 0.57671 | 0.62331 | 0.62331
| Stent | Default | 0.20096 | 0.23961 | 0.25767 | 0.28687 | 0.32038 | 0.36306 | 0.38328 | 0.38328 | 0.3617 | 0.38467 | 0.43057 | 0.4734 | 0.53445 | 0.56793 | 0.56793
Peripheral | UnSubtract | Default | 0.23267 | 0.28691 | 0.31309 | 0.35561 | 0.40488 | 0.46602 | 0.4959 | 0.4959 | 0.36495 | 0.39435 | 0.44987 | 0.47194 | 0.57671 | 0.62331 | 0.62331
| Navigate | Default | 0.23267 | 0.28691 | 0.31309 | 0.35561 | 0.40488 | 0.46602 | 0.4959 | 0.4959 | 0.36495 | 0.39435 | 0.44987 | 0.47194 | 0.57671 | 0.62331 | 0.62331

## Exampl. Settings High Air Kerma
Radioscopy | (Fluoro)
System | X-ray Protocol |  |  | Field Size | Patient Type
C12/F12 | Thorax | Subclavian | 3fps | 15 cm | Default
C15/F15 | Thorax | Subclavian | 3fps | 15 cm | Default
C20/F20 | Thorax | Subclavian | 3fps | 15 cm | Default
C12/F12 with ClarityIQ (option) | Head | Cerebral | 2fps Normal | 15 cm | Default
C15/F15 with ClarityIQ (option) | Head | Cerebral | 2fps Normal | 15 cm | Default
C20/F20 with ClarityIQ (option) | Head | Cerebral | 2fps Normal | 15 cm | Default
Radiography | Exposure
System | X-ray Protocol | Flavor | Field Size | Patient Type
C12/F12 | Head | High | 15 cm | Default
C15/F15 | Head | High | 15 cm | Default
C20/F20 | Head | High | 15 cm | Default
C12/F12 with ClarityIQ (option) | Head | Normal | 15 cm | Default
C15/F15 with ClarityIQ (option) | Head | Normal | 15 cm | Default
C20/F20 with ClarityIQ (option) | Head | Normal | 15 cm | Default